## Methods
LA-ICP-MS Trace elements data report
Laboratory & Sample Preparation
Laboratory name | FIERCE, Frankfurt Isotope & Element Research Center Goethe Univesität, Frankfurt am Main
Sample type/mineral | Garnet
Sample preparation | Polished thinsections
Imaging | Petrographic microscope & 2400 dpi digital scan
Laser ablation system
Make, Model & type | RESOlution ArF excimer laser (COMpex Pro 102)
Ablation cell | Two-volume ablation cell (Laurin Technic S155)
Laser wavelength | 193 nm
Pulse width | 20 ns
Fluence | ~ 4 J/cm2
Repetition rate | 6 Hz
Pre-ablation | 4 pulses (same parameters as main ablation)
Ablation duration | 18 s
Ablation rate | ~ 0.6 μm/s
Spot shape & size | Round 33 μm (diameter) spots
Sampling mode | Static spot ablation
Gasses | Sample cell: He. Funnel: He + Ar. Tubbing: He + Ar + N
Gas flows | He (300 ml/min), Ar (1100 ml/min), N (5 ml/min).
ICP-MS Instrument
Make, Model & type | ThermoScientific ElementXR sector field ICP-MS
Sample introduction | Ablation aerosol
RF power | 1300 W
Detection system | Secondary electron multiplier (with conversion dynode at -8kV). Simultaneous analogue and counting (pulse) modes of detection (conversion factors calculated per mass and applied offline). Magnetic field fixed. Detection by peak jumping with electrostatic analyzer.
Masses measured | 23Na, 25Mg, 27Al, 29Si, 43Ca, 44Ca, 49Ti, 51V, 52Cr, 55Mn, 57Fe, 59Co, 60Ni, 66Zn, 88Sr, 89Y, 90Zr, 93Nb, 139La, 140Ce, 141Pr, 146Nd, 147Sm, 151Eu, 157Gd, 158Gd, 159Tb, 161Dy, 165Ho, 167Er, 172Yb, 175Lu, 178Hf, 181Ta, 206Pb, 208Pb, 232Th, 238U
Dwell times | 1.2 ms for masses 23, 25, 27, 43, and 44, 1.8 Ms for all other masses
Samples per peak/integration type | 4 for all masses/average
Total time per cycle | 1.007 s
Number of cycles/total time | 44 / 43.2 s
Acquisition mode | Trigger from laser (20 s after pre-ablation), background: 20 s, ablation: 20 s
Dead time
Data Processing
Gas blank | 20 s on-peak zero subtracted.
Calibration strategy | - Primary RM: GSD-1G - Validation RM: NIST SRM612, BCR-1, BIR-1G
Reference Material (RM) information | Soda-lime glass NIST SRM612 (Jochum et al., 2011), basalt glasses GSD-1G, BCR-1, BIR-1G (Jochum et al., 2005).
Data processing / LIEF correction | Trace element data were processed using LADR (Norris & Danyushevsky, 2018).
Uncertainty level & propagation | Uncertainties are quoted at 2s absolute and are propagated in LADR. These uncertainties include those in the background signal, peak signal, and uncertainties in reference material analyses, as well as propagated systematic uncertainties in the published reference material values.

## OM20-13
Table 4b. LA-ICP-MS trace element analyses (33 μm) of garnet from sample OM20-13.
| Major elements (wt %) | [13 empty] | Trace elements (μg/g)
Analyses | Na | 2s | Mg | 2s | Al | 2s | Ca (43) | 2s | Ca (44) | 2s | Mn | 2s | Fe | 2s | Ti | 2s | V | 2s | Cr | 2s | Co | 2s | Ni | 2s | Zn | 2s | Sr | 2s | Y | 2s | Zr | 2s | Nb | 2s | La | 2s | Ce | 2s | Pr | 2s | Nd | 2s | Sm | 2s | Eu | 2s | Gd (157) | 2s | Gd (158) | 2s | Tb | 2s | Dy | 2s | Ho | 2s | Er | 2s | Tm | 2s | Yb | 2s | Lu | 2s | Hf | 2s | Ta | 2s | Pb (206) | 2s | Pb (208) | 2s | Th | 2s | U | 2s | Notes
9 | 0.00657 | 0.00014229 | 2.4266 | 0.068146 | 10.2739 | 0.26856 | 2.0113 | 0.089098 | 1.9265 | 0.0485 | 0.11928 | 0.002432 | 25.9894 | 0.58694 | 162.2843 | 3.9663 | 51.722 | 1.6312 | 71.6176 | 3.6076 | 60.7872 | 1.5905 | 5.4895 | 1.4175 | 353.7592 | 11.9527 | 0.17698 | 0.039709 | 11.4416 | 0.31909 | 7.2869 | 1.2083 | bdl |  | 0.2164 | 0.023471 | 0.1577 | 0.022182 | 0.069223 | 0.011037 | 0.44738 | 0.075081 | 1.7879 | 0.10976 | 0.98748 | 0.059497 | 4.3738 | 0.18681 | 4.2608 | 0.20115 | 0.31296 | 0.018296 | 1.4992 | 0.11486 | 0.3833 | 0.025411 | 1.635 | 0.12134 | 0.35538 | 0.021301 | 3.3168 | 0.22914 | 0.61616 | 0.029404 | 0.13405 | 0.043999 | bdl |  | 0.59143 | 0.10151 | 0.62445 | 0.051636 | 0.27051 | 0.042591 | 0.065041 | 0.016077 | Zrn contamination
10 | 0.009649 | 0.0002179 | 2.9251 | 0.13882 | 10.6451 | 0.5174 | 2.0066 | 0.090343 | 2.0319 | 0.083383 | 0.063095 | 0.0033413 | 26.0918 | 1.0878 | 117.4387 | 10.3383 | 40.9937 | 1.1571 | 66.2892 | 4.3126 | 76.3501 | 2.4093 | 4.1942 | 0.2683 | 427.649 | 14.0207 | bdl |  | 4.799 | 0.18329 | 4.5203 | 1.1421 | bdl |  | bdl |  | 0.0045878 | 0.0027538 | 0.0038755 | 0.002255 | 0.099618 | 0.028023 | 1.3065 | 0.13544 | 0.82455 | 0.10401 | 4.049 | 0.36076 | 3.8248 | 0.26002 | 0.25183 | 0.024975 | 0.7138 | 0.11386 | 0.14609 | 0.019569 | 0.66416 | 0.090402 | 0.14101 | 0.032693 | 1.0784 | 0.15008 | 0.20058 | 0.025777 | 0.24384 | 0.2856 | bdl |  | 0.041258 | 0.015655 | 0.048776 | 0.018269 | 0.029167 | 0.012026 | 0.049126 | 0.017311 | Zrn contamination
11 | 0.0085502 | 0.00022688 | 2.6638 | 0.061386 | 9.8554 | 0.33144 | 2.0419 | 0.099774 | 2.0275 | 0.16483 | 0.058053 | 0.0049925 | 24.8909 | 0.57685 | 146.8324 | 15.4399 | 46.4648 | 1.1877 | 125.3369 | 6.075 | 72.2161 | 1.7636 | 2.5527 | 2.5471 | 434.6129 | 57.9702 | bdl |  | 5.458 | 0.24888 | 1.2131 | 0.27804 | bdl |  | bdl |  | bdl |  | 0.004839 | 0.0045305 | 0.19204 | 0.059068 | 1.6791 | 0.42037 | 1.0577 | 0.095401 | 4.8946 | 0.34574 | 5.5351 | 0.50465 | 0.29744 | 0.054591 | 0.88113 | 0.13573 | 0.17012 | 0.031212 | 0.73553 | 0.10548 | 0.18467 | 0.030738 | 1.2586 | 0.22603 | 0.24709 | 0.038946 | 0.035662 | 0.020966 | bdl |  | 0.02703 | 0.018283 | 0.01951 | 0.010937 | 0.0008141 | 0.001784 | 0.01705 | 0.0067694
12 | 0.0093265 | 0.00019971 | 2.858 | 0.095386 | 10.4708 | 0.46678 | 2.1428 | 0.12499 | 2.0788 | 0.084874 | 0.059261 | 0.003012 | 25.6548 | 0.80338 | 147.2235 | 5.9321 | 49.6801 | 2.1305 | 137.9259 | 7.4934 | 76.2761 | 1.8571 | 4.722 | 0.26183 | 470.6338 | 19.2923 | bdl |  | 5.5394 | 0.24406 | 1.2509 | 0.33539 | bdl |  | bdl |  | bdl |  | 0.0038958 | 0.0020805 | 0.13367 | 0.039203 | 1.815 | 0.16331 | 1.0511 | 0.054116 | 5.6442 | 0.29351 | 5.6644 | 0.32254 | 0.34078 | 0.020002 | 0.83913 | 0.092267 | 0.17677 | 0.022739 | 0.79889 | 0.12039 | 0.16332 | 0.013959 | 1.5072 | 0.13199 | 0.26227 | 0.03031 | 0.029261 | 0.011091 | bdl |  | 0.026812 | 0.010627 | 0.026119 | 0.0073906 | 0.0075557 | 0.0031712 | 0.0154 | 0.0035095
13 | 0.0031013 | 8.6203e-05 | 2.3153 | 0.15465 | 10.7362 | 0.55131 | 1.9351 | 0.1193 | 1.8863 | 0.15179 | 0.14595 | 0.010231 | 29.244 | 3.0074 | 162.4217 | 10.6307 | 52.2766 | 6.387 | 68.1099 | 19.2331 | 54.0832 | 2.2486 | 1.5725 | 0.18075 | 450.8071 | 100.5003 | bdl |  | 14.467 | 0.58152 | 2.0611 | 0.50336 | bdl |  | bdl |  | bdl |  | bdl |  | 0.15438 | 0.069012 | 1.641 | 0.20129 | 0.90278 | 0.087313 | 3.9322 | 0.34604 | 4.1268 | 0.21569 | 0.36461 | 0.059275 | 1.7554 | 0.15225 | 0.43715 | 0.047665 | 2.2955 | 0.1908 | 0.56225 | 0.063306 | 4.8823 | 0.57531 | 0.97935 | 0.08419 | 0.023565 | 0.014083 | bdl |  | 0.051898 | 0.021362 | 0.033814 | 0.013845 | 0.011978 | 0.0081861 | 0.022756 | 0.0077995 | Zrn contamination
14 | 0.0096201 | 0.00020407 | 2.8444 | 0.11098 | 10.357 | 0.42082 | 2.0174 | 0.076232 | 1.9334 | 0.057055 | 0.057489 | 0.0022777 | 24.5013 | 0.71155 | 204.0157 | 39.8702 | 54.0496 | 1.3534 | 130.223 | 5.6548 | 75.5077 | 2.123 | 4.1073 | 0.23433 | 454.1769 | 16.7591 | 0.27778 | 0.025135 | 6.4711 | 0.18619 | 1.3023 | 0.24156 | 0.030975 | 0.0073061 | 0.27416 | 0.032327 | 0.032081 | 0.0077923 | 0.091011 | 0.013994 | 0.52309 | 0.070541 | 1.6218 | 0.13655 | 0.91827 | 0.068757 | 4.883 | 0.19974 | 4.845 | 0.2252 | 0.28437 | 0.021467 | 0.87687 | 0.077745 | 0.23021 | 0.02668 | 0.99655 | 0.074643 | 0.20324 | 0.01866 | 1.8817 | 0.15687 | 0.31338 | 0.023561 | 0.018337 | 0.0081933 | bdl |  | 0.17083 | 0.02677 | 0.16746 | 0.021298 | 0.099704 | 0.01481 | 0.021698 | 0.0040131
15 | 0.012457 | 0.00026424 | 2.8499 | 0.095588 | 10.3735 | 0.35421 | 1.9821 | 0.06268 | 1.9136 | 0.054105 | 0.063664 | 0.0020418 | 24.8676 | 0.77473 | 125.8392 | 7.6415 | 45.4812 | 1.1816 | 105.495 | 4.42 | 74.2152 | 2.5899 | 3.8801 | 0.22466 | 449.9564 | 22.8626 | 0.18648 | 0.03253 | 5.5038 | 0.16161 | 1.0363 | 0.095365 | bdl |  | 0.10442 | 0.013889 | 0.008837 | 0.0032455 | 0.036353 | 0.013904 | 0.25596 | 0.049214 | 1.5574 | 0.15534 | 0.94745 | 0.068774 | 4.2616 | 0.19461 | 4.3139 | 0.15348 | 0.28333 | 0.028079 | 0.80354 | 0.10076 | 0.15946 | 0.013326 | 0.76673 | 0.081047 | 0.15534 | 0.012422 | 1.3683 | 0.11825 | 0.22853 | 0.016035 | 0.009807 | 0.0059355 | bdl |  | 0.14029 | 0.028059 | 0.13685 | 0.02559 | 0.039461 | 0.0081212 | 0.018927 | 0.0062159
16 | 0.14795 | 0.031473 | 1.8829 | 0.07412 | 10.4858 | 0.28012 | 2.0985 | 0.066666 | 2.0873 | 0.1102 | 0.10878 | 0.0023135 | 29.1366 | 3.3526 | 271.6355 | 31.9432 | 68.2472 | 2.5619 | 71.6257 | 3.3129 | 49.3844 | 2.6146 | 3.2772 | 0.23552 | 693.4272 | 640.9958 | bdl |  | 9.2173 | 0.24851 | 3.86 | 1.5545 | 0.067616 | 0.02616 | 0.01788 | 0.005261 | 0.040709 | 0.013136 | 0.0092459 | 0.0038318 | 0.22541 | 0.053121 | 2.4464 | 0.2096 | 1.1072 | 0.088467 | 4.4091 | 0.26838 | 4.71 | 0.24243 | 0.434 | 0.025919 | 1.6413 | 0.13186 | 0.27959 | 0.023997 | 1.1402 | 0.095437 | 0.25687 | 0.025231 | 2.2718 | 0.13531 | 0.41367 | 0.03993 | 0.065217 | 0.018701 | 0.0041597 | 0.0021743 | 0.19958 | 0.051717 | 0.25546 | 0.06213 | 0.10743 | 0.024302 | 0.067825 | 0.024241 | Zrn contamination
17 | 0.60889 | 0.087721 | 2.4518 | 0.40781 | 10.5043 | 0.37376 | 1.9385 | 0.10045 | 1.9746 | 0.1682 | 0.11333 | 0.0049093 | 26.0155 | 0.92377 | 162.7145 | 26.6347 | 58.8043 | 3.4645 | 86.7969 | 4.2833 | 53.2221 | 1.8308 | 1.2269 | 0.16767 | 360.2244 | 26.2256 | bdl |  | 8.9674 | 0.2752 | 1.3475 | 0.69775 | 0.032343 | 0.013489 | 0.010515 | 0.0049774 | 0.031384 | 0.0089144 | 0.0043173 | 0.0032447 | 0.16569 | 0.044784 | 1.5002 | 0.14707 | 0.70224 | 0.096544 | 3.4397 | 0.3602 | 3.301 | 0.37558 | 0.2516 | 0.040002 | 1.1178 | 0.1173 | 0.2826 | 0.043593 | 1.3488 | 0.12562 | 0.24585 | 0.031673 | 2.6005 | 0.18093 | 0.44172 | 0.035986 | 0.025889 | 0.014468 | bdl |  | 0.099826 | 0.045181 | 0.20915 | 0.068038 | 0.10997 | 0.016839 | 0.059061 | 0.013127 | Zrn contamination
18 | 0.055077 | 0.023833 | 2.0177 | 0.063246 | 10.1975 | 0.30836 | 1.7269 | 0.045828 | 1.6056 | 0.029834 | 0.13884 | 0.0039342 | 27.1204 | 0.57276 | 126.8918 | 5.0106 | 38.5685 | 0.89875 | 66.1623 | 3.392 | 50.3349 | 1.2819 | 1.5552 | 0.1391 | 344.9145 | 13.054 | 1.8581 | 0.13706 | 12.4611 | 0.24856 | 5.1391 | 2.2427 | 0.022188 | 0.0063634 | 0.13476 | 0.024992 | 0.090152 | 0.016708 | 0.031968 | 0.0069215 | 0.23845 | 0.042291 | 1.3088 | 0.11677 | 0.69224 | 0.045787 | 3.2436 | 0.16234 | 2.9815 | 0.16934 | 0.27894 | 0.018569 | 1.571 | 0.13355 | 0.4295 | 0.025128 | 2.0025 | 0.13629 | 0.47047 | 0.026033 | 4.4833 | 0.23537 | 0.84406 | 0.042986 | 0.093019 | 0.038084 | bdl |  | 1.1028 | 0.10208 | 1.1831 | 0.11271 | 0.24594 | 0.025307 | 0.082484 | 0.019952 | Zrn contamination
19 | 0.0080602 | 0.00017332 | 1.5063 | 0.037094 | 10.1333 | 0.30393 | 2.4408 | 0.079517 | 2.3828 | 0.057374 | 0.11654 | 0.0034122 | 26.7183 | 0.46366 | 422.277 | 26.1232 | 64.5454 | 5.7882 | 87.5682 | 4.1588 | 43.4631 | 4.9258 | 3.2351 | 0.21312 | 296.7131 | 13.0788 | 0.053285 | 0.010031 | 21.6817 | 1.2926 | 965.61 | 520.2183 | 0.093449 | 0.01337 | 0.02225 | 0.0052915 | 0.029181 | 0.0062756 | 0.024882 | 0.0070932 | 0.53624 | 0.095672 | 5.2486 | 0.32663 | 2.6263 | 0.11115 | 11.3093 | 0.3932 | 11.5475 | 0.39308 | 1.0714 | 0.045701 | 4.4491 | 0.21278 | 0.68433 | 0.052966 | 2.5783 | 0.21325 | 0.5396 | 0.055755 | 5.2102 | 0.62869 | 0.98908 | 0.10408 | 24.2032 | 11.997 | 0.08927 | 0.045471 | 0.35135 | 0.13916 | 0.099612 | 0.028847 | 0.75995 | 0.33429 | 6.6303 | 3.5472 | Zrn contamination
20 | 1.4431 | 0.12221 | 1.3404 | 0.03892 | 8.7502 | 0.31053 | 2.1415 | 0.05242 | 2.1132 | 0.052971 | 0.092964 | 0.0043314 | 23.7612 | 1.0791 | 319.1549 | 34.6044 | 88.0778 | 5.2606 | 91.4564 | 4.1781 | 32.614 | 0.84823 | 2.3774 | 0.1828 | 244.7341 | 10.4367 | 0.19374 | 0.028916 | 12.1025 | 0.34845 | 3.6814 | 1.642 | 0.039237 | 0.033331 | 0.020956 | 0.0050406 | 0.053894 | 0.012721 | 0.027316 | 0.0065158 | 0.34557 | 0.060062 | 3.0083 | 0.21226 | 1.4764 | 0.086115 | 6.6792 | 0.37314 | 6.5987 | 0.2282 | 0.63502 | 0.030633 | 2.467 | 0.15218 | 0.41144 | 0.031986 | 1.4036 | 0.096235 | 0.27186 | 0.019094 | 2.8126 | 0.18755 | 0.49599 | 0.026428 | 0.090003 | 0.019434 | 0.0047937 | 0.0020632 | 0.18821 | 0.046961 | 0.17886 | 0.028309 | 0.17223 | 0.056306 | 0.14 | 0.035258 | Zrn contamination
21 | 0.0086985 | 0.00018595 | 1.3008 | 0.026673 | 10.0418 | 0.41356 | 2.472 | 0.062088 | 2.3662 | 0.047384 | 0.35452 | 0.0066287 | 27.0019 | 0.53236 | 321.7559 | 22.8653 | 29.1684 | 0.66509 | 68.2179 | 3.2028 | 24.2941 | 0.41615 | 1.9304 | 0.15951 | 219.758 | 7.8298 | bdl |  | 186.4746 | 3.0676 | 51.1811 | 25.6182 | 0.080944 | 0.01865 | bdl |  | 0.016563 | 0.0047101 | 0.01134 | 0.0034773 | 0.35814 | 0.074511 | 3.8929 | 0.27314 | 2.6208 | 0.11463 | 25.6147 | 0.69412 | 26.2201 | 0.74195 | 6.0014 | 0.16425 | 41.3822 | 1.1928 | 7.2955 | 0.20955 | 16.4253 | 0.5027 | 1.9431 | 0.067278 | 13.277 | 0.49624 | 2.1606 | 0.088446 | 1.089 | 0.54979 | 0.035408 | 0.007475 | 0.018177 | 0.0085542 | 0.019025 | 0.0061786 | 0.10224 | 0.025539 | 0.33508 | 0.13818 | Zrn contamination
22 | 0.012658 | 0.00027397 | 1.3823 | 0.043781 | 10.2233 | 0.24913 | 2.3514 | 0.086149 | 2.2835 | 0.072347 | 0.27283 | 0.013774 | 27.7129 | 0.85988 | 322.6011 | 41.0716 | 39.838 | 0.87547 | 58.5007 | 4.3827 | 29.1622 | 0.75625 | 2.5303 | 0.19295 | 238.1521 | 9.627 | bdl |  | 98.3766 | 2.2386 | 41.6153 | 13.434 | 1.4741 | 1.6187 | bdl |  | 0.0077615 | 0.0033656 | 0.0055858 | 0.0025495 | 0.23505 | 0.04559 | 2.4949 | 0.26423 | 1.8389 | 0.11047 | 16.4317 | 0.76223 | 16.3125 | 0.49446 | 3.6497 | 0.091245 | 23.0535 | 0.84584 | 3.7061 | 0.10335 | 7.8617 | 0.3257 | 0.80289 | 0.035743 | 5.1781 | 0.27416 | 0.79631 | 0.03787 | 0.86073 | 0.31667 | 0.20769 | 0.22676 | 0.14222 | 0.039729 | 0.078613 | 0.037659 | 0.24173 | 0.057199 | 0.46898 | 0.13626 | Zrn contamination
23 | 0.0057777 | 0.00012789 | 1.1999 | 0.03311 | 9.4272 | 0.30114 | 2.3529 | 0.080526 | 2.3066 | 0.069922 | 0.3621 | 0.010158 | 25.2519 | 0.66909 | 289.6553 | 11.2482 | 31.9122 | 1.0147 | 69.1666 | 3.4073 | 22.4629 | 0.49286 | 2.2293 | 0.17213 | 241.197 | 8.6984 | bdl |  | 153.238 | 4.5015 | 1.0284 | 0.14547 | 0.34567 | 0.025881 | bdl |  | bdl |  | bdl |  | 0.19915 | 0.05352 | 2.3141 | 0.18269 | 1.7664 | 0.075083 | 18.2592 | 0.54167 | 18.3915 | 0.59178 | 4.6452 | 0.15553 | 33.461 | 1.4056 | 5.8737 | 0.24547 | 13.8034 | 0.54253 | 1.4439 | 0.071296 | 10.0285 | 0.40992 | 1.586 | 0.070926 | 0.018684 | 0.0086158 | 0.033234 | 0.031151 | 0.030705 | 0.011053 | 0.024277 | 0.0069395 | 0.0017268 | 0.0014724 | 0.013276 | 0.0041032
24 | 0.0086896 | 0.00019181 | 1.1443 | 0.026695 | 10.0661 | 0.28756 | 2.846 | 0.095615 | 2.8156 | 0.056593 | 0.82541 | 0.018753 | 26.7215 | 0.47832 | 305.4495 | 10.4667 | 30.2467 | 0.6789 | 79.1572 | 3.3747 | 14.0357 | 0.31021 | 1.0147 | 0.12976 | 206.855 | 7.2963 | 0.050483 | 0.011229 | 279.8344 | 5.9315 | 2.5472 | 0.89621 | 0.54408 | 0.70075 | bdl |  | bdl |  | 0.0074447 | 0.0032453 | 0.14904 | 0.034164 | 1.9946 | 0.25401 | 1.5786 | 0.075897 | 20.6381 | 0.75591 | 20.929 | 0.52309 | 6.2755 | 0.23181 | 52.6541 | 1.9253 | 10.8079 | 0.40139 | 30.487 | 1.167 | 4.1396 | 0.15675 | 30.6981 | 0.90023 | 5.2612 | 0.17478 | 0.039394 | 0.014338 | 0.021512 | 0.0048999 | 0.022197 | 0.010758 | 0.015857 | 0.011362 | 0.0089714 | 0.003845 | 0.032454 | 0.013274 | Zrn contamination
25 | 0.0043282 | 9.8251e-05 | 1.1367 | 0.049107 | 10.0404 | 0.45667 | 2.6475 | 0.1079 | 2.5922 | 0.083935 | 0.67095 | 0.023955 | 25.8768 | 0.78696 | 300.4423 | 12.0163 | 28.9165 | 1.1285 | 85.2889 | 4.3512 | 15.1699 | 0.59763 | 0.16203 | 0.042736 | 199.2909 | 11.4965 | bdl |  | 235.4364 | 7.3343 | 119.0316 | 45.8136 | 0.084493 | 0.019262 | bdl |  | 0.010742 | 0.0036632 | 0.0073319 | 0.0044584 | 0.17793 | 0.033632 | 2.1142 | 0.13814 | 1.6865 | 0.10166 | 20.0447 | 0.79902 | 20.5038 | 0.81952 | 5.7279 | 0.24056 | 46.4831 | 2.2394 | 9.1055 | 0.27675 | 24.4177 | 0.96167 | 3.1317 | 0.10808 | 22.9807 | 0.95665 | 3.9182 | 0.17824 | 2.9285 | 1.1348 | 0.057856 | 0.022632 | 0.095764 | 0.018917 | 0.032463 | 0.011204 | 0.28622 | 0.092498 | 0.88876 | 0.3039 | Zrn contamination
26 | 0.0070469 | 0.00016068 | 1.1835 | 0.039415 | 10.4156 | 0.34917 | 2.7405 | 0.11517 | 2.7317 | 0.091451 | 0.85068 | 0.026819 | 27.9628 | 0.70432 | 317.3929 | 11.439 | 29.7144 | 0.97489 | 101.0006 | 5.2594 | 14.3007 | 0.38493 | 0.073059 | 0.035433 | 201.3768 | 7.6947 | bdl |  | 270.3296 | 9.388 | 9.0551 | 2.7524 | 0.23842 | 0.35067 | bdl |  | bdl |  | 0.0045012 | 0.0026133 | 0.16382 | 0.038772 | 1.9675 | 0.20014 | 1.5815 | 0.08369 | 20.0137 | 0.94002 | 20.7415 | 1.3706 | 6.2109 | 0.28673 | 52.668 | 2.0904 | 10.5621 | 0.50081 | 28.9104 | 1.2975 | 3.9022 | 0.19505 | 28.2229 | 1.7777 | 4.7744 | 0.19403 | 0.32179 | 0.10798 | 0.060426 | 0.056239 | 0.022431 | 0.011151 | 0.02008 | 0.0074518 | 0.028541 | 0.011484 | 0.061667 | 0.019306 | Zrn contamination

## OM20-17
Table 4c. LA-ICP-MS trace element analyses (33 μm) of garnet from sample OM20-17.
| Major elements (wt %) | [13 empty] | Trace elements (μg/g)
Analyses | Na | 2s | Mg | 2s | Al | 2s | Ca (43) | 2s | Ca (44) | 2s | Mn | 2s | Fe | 2s | Ti | 2s | V | 2s | Cr | 2s | Co | 2s | Ni | 2s | Zn | 2s | Sr | 2s | Y | 2s | Zr | 2s | Nb | 2s | La | 2s | Ce | 2s | Pr | 2s | Nd | 2s | Sm | 2s | Eu | 2s | Gd (157) | 2s | Gd (158) | 2s | Tb | 2s | Dy | 2s | Ho | 2s | Er | 2s | Tm | 2s | Yb | 2s | Lu | 2s | Hf | 2s | Ta | 2s | Pb (206) | 2s | Pb (208) | 2s | Th | 2s | U | 2s | Notes
43 | 0.013266 | 0.0055629 | 3.0574 | 0.07232 | 10.997 | 0.25861 | 2.4024 | 0.090536 | 2.3523 | 0.054273 | 0.52681 | 0.012449 | 25.6041 | 0.45232 | 136.1824 | 4.678 | 38.9944 | 0.87415 | 22.6888 | 1.4907 | 117.2515 | 2.3196 | 5.013 | 1.1222 | 400.0869 | 12.5232 | bdl |  | 148.9008 | 3.0737 | 0.7429 | 0.1041 | bdl |  | bdl |  | 0.0049494 | 0.0033821 | 0.00084696 | 0.0012529 | 0.024964 | 0.01572 | 0.29773 | 0.061534 | 0.36257 | 0.036966 | 4.0823 | 0.23635 | 4.2876 | 0.31758 | 2.3257 | 0.082616 | 25.499 | 0.848 | 5.385 | 0.15947 | 14.6376 | 0.57752 | 1.7828 | 0.099737 | 11.0593 | 0.42248 | 1.4845 | 0.06479 | bdl |  | bdl |  | bdl |  | bdl |  | bdl |  | 0.00098534 | 0.0011764
44 | 0.024025 | 0.013876 | 3.1525 | 0.075458 | 10.7992 | 0.3295 | 2.3291 | 0.085582 | 2.2571 | 0.05496 | 0.51456 | 0.013228 | 25.1201 | 1.0976 | 123.206 | 5.1703 | 36.9267 | 1.1708 | 19.2138 | 1.6697 | 124.7228 | 2.035 | 3.1409 | 0.21988 | 428.7283 | 16.0452 | bdl |  | 140.8868 | 3.581 | 0.77613 | 0.36561 | bdl |  | bdl |  | bdl |  | bdl |  | 0.014475 | 0.0099696 | 0.28648 | 0.055807 | 0.31556 | 0.048019 | 3.7173 | 0.26407 | 3.8207 | 0.19734 | 2.2279 | 0.094335 | 24.729 | 0.92402 | 5.2929 | 0.1502 | 13.737 | 0.51451 | 1.6162 | 0.064844 | 9.9371 | 0.5054 | 1.2421 | 0.058223 | 0.015426 | 0.008376 | bdl |  | 0.014575 | 0.0081619 | bdl |  | 0.0005692 | 0.00090492 | 0.0010195 | 0.00094198
45 | 0.0096398 | 0.00022709 | 3.2889 | 0.065452 | 10.8135 | 0.25561 | 2.3071 | 0.10192 | 2.2218 | 0.053964 | 0.5033 | 0.010679 | 24.4622 | 0.43214 | 133.8079 | 7.0663 | 36.7856 | 0.78493 | 16.743 | 1.4491 | 130.9861 | 2.4262 | 4.9106 | 0.35202 | 452.5598 | 16.7304 | bdl |  | 146.546 | 4.1221 | 0.75338 | 0.10596 | bdl |  | bdl |  | bdl |  | 0.00038657 | 0.00091151 | 0.033311 | 0.019397 | 0.27571 | 0.10245 | 0.42984 | 0.051764 | 4.287 | 0.34034 | 4.4354 | 0.30003 | 2.4542 | 0.082522 | 26.6476 | 1.1036 | 5.5757 | 0.17746 | 14.7515 | 0.50705 | 1.7905 | 0.093584 | 10.4913 | 0.42213 | 1.4376 | 0.074386 | 0.013516 | 0.010038 | bdl |  | 0.15659 | 0.034379 | bdl |  | bdl |  | 0.0028 | 0.0019986
46 | 0.021643 | 0.0072813 | 2.0604 | 0.079176 | 10.9412 | 0.38952 | 3.0206 | 0.1111 | 3.0089 | 0.15091 | 1.4687 | 0.043802 | 28.9382 | 5.3298 | 234.1931 | 8.9154 | 58.517 | 2.1646 | 16.1955 | 1.5033 | 73.8998 | 1.9348 | 6.3128 | 1.3767 | 351.8082 | 13.9086 | 0.053197 | 0.029656 | 225.7911 | 7.5535 | 0.91597 | 0.42915 | 0.02004 | 0.0071094 | bdl |  | bdl |  | bdl |  | 0.017968 | 0.012351 | 0.20231 | 0.048879 | 0.34225 | 0.041751 | 4.2441 | 0.35148 | 4.2908 | 0.24829 | 2.6002 | 0.12508 | 35.0423 | 1.312 | 8.1684 | 0.31995 | 23.0445 | 0.70749 | 2.7838 | 0.083473 | 18.0583 | 0.79897 | 2.5351 | 0.13283 | 0.025901 | 0.011507 | bdl |  | 0.013634 | 0.0083646 | 0.015856 | 0.006375 | 0.0039246 | 0.002517 | 0.0028811 | 0.0016773 | Zrn contamination
47 | 0.01948 | 0.005249 | 1.9697 | 0.056804 | 10.631 | 0.26929 | 2.8107 | 0.081068 | 2.8004 | 0.07348 | 1.6482 | 0.042923 | 25.4958 | 0.48203 | 226.2043 | 7.6107 | 53.8209 | 1.3438 | 16.3432 | 1.6088 | 71.7792 | 1.2224 | 4.1994 | 0.96156 | 349.0637 | 10.6307 | 0.050128 | 0.0092142 | 229.629 | 5.1568 | 18.9535 | 12.1162 | 0.015379 | 0.0051236 | bdl |  | bdl |  | 0.00027099 | 0.00050958 | 0.012846 | 0.0083155 | 0.26459 | 0.050143 | 0.33125 | 0.035532 | 3.7533 | 0.18984 | 3.9289 | 0.16107 | 2.5081 | 0.08586 | 33.8602 | 1.2967 | 8.3287 | 0.28482 | 23.6452 | 0.81563 | 2.9371 | 0.0889 | 18.4363 | 0.84849 | 2.5509 | 0.12327 | 0.43089 | 0.040021 | bdl |  | 0.022163 | 0.0089095 | 0.020794 | 0.0083639 | 0.027118 | 0.010882 | 0.043204 | 0.019723 | Zrn contamination
48 | 0.017468 | 0.00037054 | 3.3568 | 0.057269 | 10.6751 | 0.22983 | 2.2568 | 0.060166 | 2.1596 | 0.056645 | 0.53465 | 0.0099335 | 24.4631 | 0.46618 | 88.5213 | 3.415 | 30.9505 | 0.70564 | 29.0762 | 1.6453 | 133.4291 | 2.2851 | 4.6189 | 1.4616 | 451.2172 | 15.2478 | 0.063489 | 0.010625 | 132.8731 | 2.3934 | 49.7757 | 65.3716 | bdl |  | bdl |  | bdl |  | 0.0030421 | 0.0017489 | 0.01779 | 0.010024 | 0.25126 | 0.056512 | 0.29971 | 0.032685 | 3.7297 | 0.17442 | 3.6955 | 0.1607 | 2.0556 | 0.056416 | 22.5439 | 0.68978 | 4.8841 | 0.16665 | 13.1924 | 0.43464 | 1.5592 | 0.069986 | 9.7333 | 0.42446 | 1.268 | 0.078864 | 1.2908 | 0.07379 | bdl |  | 0.012902 | 0.00696 | 0.015681 | 0.0054231 | 0.029941 | 0.0059579 | 0.095878 | 0.0084026 | Zrn contamination
49 | 0.012645 | 0.00026824 | 3.2624 | 0.060877 | 10.4779 | 0.25741 | 2.3445 | 0.061717 | 2.2167 | 0.057165 | 0.4959 | 0.0092135 | 24.3735 | 0.39356 | 115.738 | 7.263 | 36.5254 | 0.89529 | 28.2258 | 2.0049 | 131.2105 | 2.0246 | 4.4477 | 0.25205 | 433.3686 | 13.2702 | 0.31432 | 0.062602 | 139.5092 | 3.3026 | 1.1473 | 0.29334 | bdl |  | 0.0073551 | 0.0029093 | 0.016177 | 0.0044529 | 0.0046709 | 0.0021413 | 0.049318 | 0.0165 | 0.26905 | 0.041769 | 0.3574 | 0.029803 | 4.1442 | 0.28908 | 4.1578 | 0.25083 | 2.2153 | 0.086342 | 24.1779 | 0.91241 | 5.1343 | 0.17651 | 13.6844 | 0.52447 | 1.6367 | 0.055605 | 10.2219 | 0.44472 | 1.388 | 0.058807 | 0.045137 | 0.012831 | bdl |  | 0.044936 | 0.012842 | 0.036934 | 0.013317 | 0.020749 | 0.0048924 | 0.0092632 | 0.002622 | Zrn contamination
50 | 0.019929 | 0.010779 | 3.4689 | 0.10139 | 10.6066 | 0.4313 | 2.0868 | 0.086041 | 2.0684 | 0.08656 | 0.54644 | 0.013465 | 24.3107 | 0.94489 | 89.1775 | 3.5863 | 30.0352 | 1.2321 | 24.7649 | 1.5559 | 140.4809 | 3.0385 | 5.5004 | 0.2845 | 431.3736 | 18.1224 | bdl |  | 128.5285 | 3.2683 | 2.5178 | 0.68404 | bdl |  | bdl |  | bdl |  | bdl |  | 0.019772 | 0.010319 | 0.22078 | 0.050294 | 0.28972 | 0.037145 | 3.3761 | 0.16145 | 3.5455 | 0.20199 | 1.9188 | 0.075042 | 21.9711 | 0.8776 | 4.8683 | 0.18391 | 12.9943 | 0.50339 | 1.6238 | 0.077464 | 10.3346 | 0.47739 | 1.3433 | 0.060585 | 0.046701 | 0.020642 | bdl |  | 0.010607 | 0.0061533 | 0.024361 | 0.0095206 | 0.00041684 | 0.00068369 | 0.007801 | 0.0023033 | Zrn contamination
59 | 0.020475 | 0.006527 | 1.5079 | 0.12353 | 10.4274 | 0.34728 | 3.0349 | 0.26021 | 2.9364 | 0.13554 | 2.8712 | 0.078864 | 25.4158 | 0.89805 | 222.007 | 12.9489 | 57.2336 | 3.2109 | 16.8777 | 1.4947 | 53.6796 | 2.0027 | 6.628 | 1.4147 | 315.6453 | 15.4687 | 0.043872 | 0.01181 | 253.1285 | 5.0899 | 0.95754 | 0.31467 | 0.011007 | 0.0059532 | bdl |  | bdl |  | 0.0040025 | 0.0026903 | 0.010761 | 0.010789 | 0.17973 | 0.047581 | 0.21716 | 0.047989 | 2.6937 | 0.24122 | 2.848 | 0.23819 | 2.0246 | 0.071849 | 32.0195 | 1.2286 | 9.1475 | 0.34276 | 27.6666 | 1.1742 | 3.5762 | 0.10644 | 24.0107 | 0.71982 | 3.2611 | 0.12263 | 0.016159 | 0.010741 | 0.0048452 | 0.0027311 | 0.027964 | 0.014215 | 0.022537 | 0.009011 | 0.0028843 | 0.0024032 | 0.0064124 | 0.0029473 | Zrn contamination
60 | 0.022767 | 0.0033144 | 1.5912 | 0.045733 | 10.4358 | 0.3616 | 2.9536 | 0.13959 | 2.8351 | 0.10776 | 2.1294 | 0.054599 | 25.5471 | 0.966 | 242.7584 | 8.9469 | 55.9123 | 1.3168 | 16.0544 | 1.6036 | 58.2706 | 1.9968 | 5.7678 | 2.2959 | 305.1365 | 11.8357 | bdl |  | 225.2972 | 6.3026 | 0.71644 | 0.13555 | bdl |  | bdl |  | bdl |  | bdl |  | 0.016807 | 0.012061 | 0.24988 | 0.076385 | 0.3172 | 0.037952 | 3.7455 | 0.46297 | 3.8644 | 0.18009 | 2.4917 | 0.080007 | 33.5072 | 1.1847 | 8.1269 | 0.23728 | 22.834 | 0.98427 | 2.8212 | 0.15289 | 18.702 | 1.1735 | 2.6606 | 0.13012 | 0.011769 | 0.0081832 | bdl |  | 0.012439 | 0.008062 | 0.013739 | 0.005982 | 0.0023328 | 0.0019547 | 0.0021256 | 0.001523

## OM20-18
Table 4d. LA-ICP-MS trace element analyses (33 μm) of garnet from sample OM20-18.
| Major elements (wt %) | [13 empty] | Trace elements (μg/g)
Analyses | Na | 2s | Mg | 2s | Al | 2s | Ca (43) | 2s | Ca (44) | 2s | Mn | 2s | Fe | 2s | Ti | 2s | V | 2s | Cr | 2s | Co | 2s | Ni | 2s | Zn | 2s | Sr | 2s | Y | 2s | Zr | 2s | Nb | 2s | La | 2s | Ce | 2s | Pr | 2s | Nd | 2s | Sm | 2s | Eu | 2s | Gd (157) | 2s | Gd (158) | 2s | Tb | 2s | Dy | 2s | Ho | 2s | Er | 2s | Tm | 2s | Yb | 2s | Lu | 2s | Hf | 2s | Ta | 2s | Pb (206) | 2s | Pb (208) | 2s | Th | 2s | U | 2s | Notes
27 | 0.010508 | 0.0002229 | 2.8494 | 0.066969 | 10.3935 | 0.30521 | 3.0999 | 0.088266 | 3.0822 | 0.0925 | 0.22788 | 0.0051037 | 23.6133 | 0.4124 | 163.7361 | 3.8111 | 138.8495 | 3.2398 | 48.2611 | 2.6796 | 66.4517 | 1.1764 | bdl |  | 142.2118 | 4.625 | bdl |  | 47.1747 | 1.2862 | 0.53067 | 0.063591 | bdl |  | bdl |  | bdl |  | 0.0043782 | 0.0020485 | 0.19615 | 0.04275 | 1.6435 | 0.16494 | 0.98539 | 0.07367 | 10.2399 | 0.48721 | 10.376 | 0.42957 | 2.4781 | 0.090476 | 12.6607 | 0.46764 | 1.419 | 0.068061 | 2.306 | 0.14587 | 0.22696 | 0.025945 | 1.4474 | 0.15374 | 0.19747 | 0.021602 | bdl |  | bdl |  | bdl |  | bdl |  | bdl |  | 0.0058255 | 0.0020647
28 | 0.010795 | 0.00022898 | 2.6446 | 0.074993 | 10.4896 | 0.31255 | 3.0158 | 0.10015 | 3.0098 | 0.09923 | 0.32993 | 0.0093364 | 24.4028 | 0.56018 | 132.0093 | 4.4794 | 130.1882 | 3.3855 | 37.7252 | 1.8095 | 64.4045 | 1.5679 | bdl |  | 137.8791 | 5.1125 | bdl |  | 66.1723 | 1.6043 | 0.36269 | 0.034749 | bdl |  | bdl |  | bdl |  | 0.0046143 | 0.002104 | 0.16914 | 0.035338 | 1.3652 | 0.098132 | 0.91292 | 0.065227 | 9.2167 | 0.30221 | 9.4715 | 0.34047 | 2.1899 | 0.074273 | 13.7158 | 0.58581 | 2.102 | 0.071362 | 4.221 | 0.20511 | 0.39512 | 0.02873 | 2.1761 | 0.15514 | 0.19651 | 0.017113 | bdl |  | bdl |  | bdl |  | 0.0065696 | 0.003436 | bdl |  | 0.006021 | 0.002101
29 | 0.0051601 | 0.00011388 | 2.3082 | 0.074337 | 10.3508 | 0.27289 | 2.7491 | 0.10011 | 2.701 | 0.079864 | 0.2528 | 0.0098608 | 25.6897 | 0.69192 | 174.7484 | 3.9335 | 138.0544 | 3.6115 | 33.7182 | 1.6462 | 61.9269 | 1.7025 | bdl |  | 131.9955 | 5.5518 | bdl |  | 47.3657 | 1.3042 | 0.32266 | 0.052193 | bdl |  | bdl |  | 0.0057055 | 0.0025462 | 0.0061893 | 0.0023709 | 0.18177 | 0.036916 | 1.5919 | 0.11477 | 0.99087 | 0.07748 | 7.6378 | 0.36707 | 7.8221 | 0.26815 | 1.518 | 0.045566 | 9.3361 | 0.43249 | 1.5283 | 0.059077 | 3.5131 | 0.21233 | 0.37726 | 0.026926 | 1.9785 | 0.17605 | 0.22606 | 0.023766 | bdl |  | bdl |  | bdl |  | 0.010068 | 0.0041399 | bdl |  | 0.0020194 | 0.0011808
30 | 0.0097737 | 0.00020732 | 1.9181 | 0.095834 | 10.7629 | 0.57656 | 3.1977 | 0.12842 | 3.2156 | 0.088143 | 0.22236 | 0.0072138 | 26.2217 | 0.80012 | 231.7887 | 10.8896 | 125.7509 | 4.2028 | 24.54 | 1.5598 | 56.6826 | 1.5077 | bdl |  | 125.0031 | 5.375 | bdl |  | 254.7319 | 7.9227 | 0.37682 | 0.05224 | bdl |  | bdl |  | bdl |  | 0.0053604 | 0.0021804 | 0.12517 | 0.02515 | 1.5199 | 0.17548 | 1.2126 | 0.056631 | 11.3037 | 0.65722 | 11.5893 | 0.45952 | 2.9774 | 0.1366 | 30.1966 | 1.0135 | 9.0479 | 0.38785 | 37.6849 | 1.4736 | 6.2046 | 0.22414 | 46.805 | 1.9958 | 6.849 | 0.23993 | bdl |  | bdl |  | bdl |  | 0.013193 | 0.0046881 | 0.0006525 | 0.00082875 | 0.0053889 | 0.00191
31 | 0.0093959 | 0.00019931 | 1.5401 | 0.037553 | 10.9011 | 0.38118 | 3.3934 | 0.082007 | 3.3806 | 0.13804 | 0.23031 | 0.0087384 | 28.2578 | 0.7264 | 322.167 | 9.0455 | 227.9145 | 12.993 | 38.8357 | 3.5808 | 67.0726 | 3.6259 | bdl |  | 124.1741 | 4.8133 | bdl |  | 145.7058 | 3.1745 | 0.61049 | 0.051742 | bdl |  | bdl |  | bdl |  | 0.004472 | 0.001996 | 0.12532 | 0.035479 | 1.6181 | 0.13986 | 1.4005 | 0.059542 | 13.0824 | 0.49551 | 13.0912 | 0.4489 | 3.1915 | 0.081646 | 23.834 | 0.79079 | 4.7812 | 0.15258 | 13.9036 | 0.41532 | 1.8324 | 0.082512 | 11.9781 | 0.49485 | 1.5665 | 0.066559 | 0.0097851 | 0.0057124 | bdl |  | 0.014919 | 0.0070583 | 0.012349 | 0.0045373 | 0.00038721 | 0.00063856 | 0.0062691 | 0.0020649
32 | 0.0072564 | 0.00015566 | 1.3955 | 0.041192 | 10.6633 | 0.26631 | 3.4755 | 0.11789 | 3.5351 | 0.12062 | 0.20573 | 0.0083696 | 28.3678 | 1.3085 | 362.5276 | 13.9372 | 149.502 | 5.0903 | 66.271 | 3.0058 | 65.5073 | 1.2627 | bdl |  | 117.47 | 5.2776 | bdl |  | 175.238 | 3.9734 | 0.6848 | 0.052342 | bdl |  | bdl |  | bdl |  | bdl |  | 0.14237 | 0.041122 | 1.7589 | 0.11038 | 1.5955 | 0.087892 | 14.3824 | 0.54608 | 14.5214 | 0.59565 | 3.3594 | 0.08245 | 26.1311 | 0.85738 | 5.9371 | 0.21588 | 18.9843 | 0.80127 | 2.3313 | 0.10315 | 13.0547 | 0.6462 | 1.4658 | 0.060516 | 0.011831 | 0.0063823 | bdl |  | 0.01012 | 0.0059061 | bdl |  | 0.0012238 | 0.0011538 | 0.0087526 | 0.0024773
33 | 0.0054165 | 0.00011855 | 1.3877 | 0.062227 | 10.412 | 0.50715 | 3.6598 | 0.17333 | 3.8034 | 0.17656 | 0.43577 | 0.019385 | 27.8032 | 0.87279 | 373.3493 | 15.0341 | 113.5219 | 12.0205 | 33.6071 | 2.3846 | 49.3792 | 2.325 | bdl |  | 123.9335 | 6.3695 | bdl |  | 151.8224 | 3.3294 | 0.63413 | 0.055023 | 0.072338 | 0.013914 | bdl |  | bdl |  | 0.004815 | 0.0020557 | 0.14835 | 0.044884 | 2.0756 | 0.11396 | 1.6501 | 0.089232 | 16.7955 | 0.61722 | 16.8664 | 0.62874 | 3.8741 | 0.15105 | 25.3605 | 1.0146 | 4.9237 | 0.19875 | 12.9836 | 0.48412 | 1.631 | 0.062573 | 9.3353 | 0.48697 | 1.1569 | 0.063015 | bdl |  | 0.021407 | 0.0045102 | 0.011129 | 0.0060167 | 0.010377 | 0.0041066 | bdl |  | 0.012084 | 0.0040234
34 | 0.0076823 | 0.00016406 | 1.2639 | 0.047576 | 10.3131 | 0.33808 | 3.812 | 0.12468 | 3.8632 | 0.12235 | 1.672 | 0.052066 | 26.3723 | 0.73765 | 474.1582 | 15.5211 | 107.8644 | 3.1688 | 29.599 | 2.1272 | 42.1314 | 1.148 | 0.44703 | 0.069315 | 110.8227 | 3.7159 | bdl |  | 201.5559 | 6.3556 | 0.86046 | 0.073782 | 0.19777 | 0.027796 | bdl |  | bdl |  | 0.0052304 | 0.0021876 | 0.23423 | 0.042861 | 2.4477 | 0.22109 | 2.068 | 0.10719 | 20.3236 | 0.5548 | 21.1578 | 0.93885 | 4.8587 | 0.20405 | 34.1247 | 1.435 | 6.5873 | 0.26712 | 17.8824 | 0.74594 | 2.0803 | 0.072588 | 11.866 | 0.45592 | 1.4101 | 0.049581 | 0.014343 | 0.00701 | 0.035348 | 0.0077788 | 0.01168 | 0.0063294 | 0.0079027 | 0.0036786 | 0.00039836 | 0.00065639 | 0.024697 | 0.0065962
35 | 0.011782 | 0.00024992 | 1.2597 | 0.050515 | 10.4773 | 0.373 | 3.7693 | 0.13933 | 3.8349 | 0.16777 | 2.0489 | 0.071677 | 26.0739 | 0.89452 | 485.459 | 21.2414 | 114.0885 | 7.6397 | 31.2563 | 2.3108 | 38.7977 | 1.0662 | bdl |  | 110.8269 | 5.1324 | bdl |  | 206.2753 | 5.2939 | 0.83282 | 0.077971 | 0.25398 | 0.034192 | bdl |  | 0.0071116 | 0.0028426 | 0.0079284 | 0.0026853 | 0.23575 | 0.044957 | 2.4593 | 0.14175 | 2.085 | 0.1104 | 20.3063 | 0.70821 | 20.6719 | 0.5972 | 4.9681 | 0.13515 | 34.315 | 1.2864 | 6.8264 | 0.1965 | 18.902 | 0.67494 | 2.2864 | 0.085347 | 12.946 | 0.48034 | 1.6001 | 0.064195 | bdl |  | 0.038685 | 0.0081779 | 0.015891 | 0.0073689 | 0.012589 | 0.004635 | bdl |  | 0.024318 | 0.0047535
36 | 0.0079541 | 0.00016902 | 1.2161 | 0.070352 | 10.0239 | 0.31649 | 3.3257 | 0.10374 | 3.4125 | 0.18437 | 1.8429 | 0.060182 | 25.771 | 0.96855 | 351.8928 | 13.1598 | 258.5647 | 15.3753 | 16.7787 | 1.9141 | 36.6174 | 1.0915 | 0.60164 | 0.079439 | 105.4149 | 7.6437 | bdl |  | 118.037 | 4.1666 | 0.48605 | 0.041912 | 0.037795 | 0.015244 | bdl |  | bdl |  | 0.0043324 | 0.0019549 | 0.15811 | 0.028209 | 1.4643 | 0.10954 | 1.3468 | 0.087476 | 13.2522 | 0.4962 | 13.6024 | 0.67908 | 2.9455 | 0.11089 | 19.9162 | 0.8184 | 4.0292 | 0.15833 | 11.7274 | 0.43141 | 1.4736 | 0.096302 | 8.6035 | 0.34952 | 1.0014 | 0.066278 | bdl |  | 0.0035605 | 0.0015869 | 0.01535 | 0.0071267 | 0.020095 | 0.005766 | 0.0011741 | 0.0011065 | 0.0071906 | 0.0024402
37 | 0.01224 | 0.00025965 | 1.3709 | 0.051729 | 10.5591 | 0.35113 | 3.3574 | 0.10761 | 3.3148 | 0.12714 | 1.8881 | 0.061723 | 26.1172 | 0.65134 | 281.7565 | 11.3955 | 168.7074 | 4.1315 | 16.9235 | 1.2658 | 40.9979 | 1.0602 | 2.6047 | 0.17937 | 112.4711 | 4.2327 | bdl |  | 137.3231 | 4.2296 | 0.53676 | 0.040136 | 0.067208 | 0.018326 | bdl |  | bdl |  | 0.0050003 | 0.0021539 | 0.11352 | 0.036111 | 1.5789 | 0.2012 | 1.3953 | 0.073923 | 13.4221 | 0.6857 | 14.281 | 0.53142 | 3.1922 | 0.14344 | 22.2485 | 0.97098 | 4.6447 | 0.20026 | 13.8306 | 0.67859 | 1.7762 | 0.072967 | 10.7868 | 0.51584 | 1.3151 | 0.061064 | bdl |  | 0.01217 | 0.0039781 | bdl |  | 0.011921 | 0.004551 | 0.00095785 | 0.0010247 | 0.0096774 | 0.0026111
38 | 0.014159 | 0.00030034 | 1.2267 | 0.044481 | 10.3938 | 0.30761 | 3.4476 | 0.12162 | 3.3197 | 0.085934 | 5.4954 | 0.12442 | 23.0454 | 0.63193 | 412.0722 | 19.2117 | 109.8966 | 12.9584 | 26.812 | 2.047 | 35.6464 | 5.6641 | 1.6227 | 0.14133 | 101.8293 | 4.2234 | bdl |  | 702.6837 | 24.0159 | 0.48676 | 0.060088 | 0.082509 | 0.015859 | bdl |  | bdl |  | bdl |  | 0.14311 | 0.038512 | 1.9441 | 0.18255 | 1.6699 | 0.12407 | 30.5211 | 0.91469 | 31.2707 | 1.0491 | 14.6129 | 0.39774 | 131.6581 | 4.8581 | 22.4198 | 0.69805 | 46.5528 | 1.378 | 4.4567 | 0.16388 | 23.0727 | 1.024 | 2.8635 | 0.13504 | 0.0087916 | 0.0056655 | 0.031775 | 0.0063668 | bdl |  | bdl |  | 0.00042636 | 0.0007017 | 0.0098343 | 0.0031063
39 | 0.0059692 | 0.00013085 | 1.1032 | 0.028459 | 10.3898 | 0.26303 | 3.6379 | 0.13436 | 3.6737 | 0.15306 | 5.4375 | 0.1295 | 22.8912 | 0.51644 | 473.3949 | 20.937 | 122.2787 | 2.4998 | 41.7953 | 3.9461 | 30.4808 | 0.56495 | 1.8985 | 0.15345 | 100.9478 | 3.4352 | 0.057349 | 0.0099266 | 538.6002 | 36.577 | 0.51344 | 0.05412 | 0.074326 | 0.017013 | bdl |  | 0.0081064 | 0.0031348 | 0.0073815 | 0.0026773 | 0.22754 | 0.046025 | 2.2171 | 0.1844 | 1.842 | 0.11089 | 30.7742 | 1.5268 | 31.0047 | 1.6527 | 13.2323 | 0.66608 | 106.5623 | 6.144 | 16.9072 | 1.1191 | 34.705 | 2.6516 | 3.4861 | 0.25818 | 18.7068 | 1.5435 | 2.3595 | 0.25938 | 0.011432 | 0.0064325 | 0.023655 | 0.0057032 | 0.013121 | 0.0069021 | 0.014254 | 0.0050853 | 0.0061963 | 0.0026632 | 0.012739 | 0.0032012
40 | 0.019874 | 0.0038909 | 1.8625 | 0.062321 | 10.4594 | 0.50853 | 3.2601 | 0.1941 | 3.2851 | 0.17326 | 2.9441 | 0.14253 | 25.272 | 0.72484 | 416.1198 | 18.5828 | 118.2126 | 3.8521 | 26.1203 | 2.2363 | 51.1423 | 1.4362 | 2.9082 | 0.19074 | 100.7239 | 4.4865 | 0.046108 | 0.0086657 | 662.5962 | 37.4145 | 0.49105 | 0.042037 | 0.095764 | 0.017244 | bdl |  | 0.0091753 | 0.00325 | 0.0055864 | 0.0022694 | 0.25643 | 0.044167 | 2.2368 | 0.14342 | 1.6846 | 0.10356 | 24.3818 | 1.222 | 24.8172 | 1.1037 | 11.0759 | 0.46978 | 109.3836 | 6.3637 | 20.2709 | 1.3107 | 42.3752 | 2.9687 | 3.9272 | 0.20603 | 19.9568 | 0.94742 | 2.267 | 0.10365 | bdl |  | 0.028787 | 0.0062649 | bdl |  | 0.020923 | 0.0060088 | 0.0012245 | 0.0011536 | 0.020666 | 0.0038063
41 | 0.023628 | 0.0097303 | 1.4711 | 0.044471 | 10.9645 | 0.59537 | 3.609 | 0.12851 | 3.5493 | 0.13544 | 6.3924 | 0.18223 | 21.6685 | 0.4062 | 469.052 | 17.7129 | 175.3189 | 4.1956 | 33.2951 | 2.0438 | 39.4314 | 0.84333 | 1.2664 | 0.13759 | 98.6676 | 4.7626 | 0.14557 | 0.017836 | 522.7133 | 17.7672 | 0.56025 | 0.075408 | 0.16991 | 0.020882 | 0.031326 | 0.0067261 | 0.017805 | 0.0076317 | 0.012965 | 0.0039885 | 0.19765 | 0.040217 | 1.494 | 0.12907 | 1.3359 | 0.11193 | 17.0414 | 0.82993 | 17.3252 | 0.88872 | 7.9848 | 0.25701 | 78.9889 | 3.216 | 15.5135 | 0.57415 | 38.6024 | 1.1467 | 4.5844 | 0.15407 | 26.7761 | 0.94109 | 3.3252 | 0.1327 | bdl |  | 0.041806 | 0.0090066 | 0.053061 | 0.015504 | 0.049672 | 0.024021 | 0.033967 | 0.0072386 | 0.027352 | 0.0073424
42 | 0.013799 | 0.0002927 | 1.5309 | 0.032757 | 10.2482 | 0.30257 | 3.5823 | 0.095735 | 3.5336 | 0.10964 | 5.7223 | 0.16141 | 21.3852 | 0.39899 | 567.4195 | 20.9104 | 165.4967 | 4.6061 | 38.3845 | 2.1539 | 41.9975 | 0.9268 | 5.0557 | 0.27096 | 94.3816 | 4.0037 | 0.045669 | 0.0088475 | 656.0562 | 18.0211 | 0.64055 | 0.052243 | 1.0087 | 0.9711 | bdl |  | 0.0054189 | 0.0025602 | 0.0077742 | 0.0027463 | 0.18622 | 0.059788 | 1.5514 | 0.14578 | 1.4206 | 0.082242 | 19.5851 | 0.74136 | 19.5835 | 0.5909 | 9.4516 | 0.25447 | 96.8748 | 3.3001 | 19.1851 | 0.62059 | 45.4668 | 1.3532 | 4.8752 | 0.17295 | 27.2612 | 1.0012 | 3.1925 | 0.10968 | 0.023181 | 0.0091555 | 0.15761 | 0.11057 | bdl |  | 0.013307 | 0.0049098 | 0.0015749 | 0.0013409 | 0.021055 | 0.0071444

## NIST SRM614
Table 4e. LA-ICP-MS trace element analyses (33 μm) of secondary reference glass NIST SRM612.
| Major elements | [13 empty] | Trace elements (μg/g)
Analyses | Na (wt %) | 2s | Mg (μg/g) | 2s | Al (wt %) | 2s | Ca (43) (wt %) | 2s | Ca (44) (wt %) | 2s | Mn (μg/g) | 2s | Fe (μg/g) | 2s | Ti | 2s | V | 2s | Cr | 2s | Co | 2s | Ni | 2s | Zn | 2s | Sr | 2s | Y | 2s | Zr | 2s | Nb | 2s | La | 2s | Ce | 2s | Pr | 2s | Nd | 2s | Sm | 2s | Eu | 2s | Gd (157) | 2s | Gd (158) | 2s | Tb | 2s | Dy | 2s | Ho | 2s | Er | 2s | Tm | 2s | Yb | 2s | Lu | 2s | Hf | 2s | Ta | 2s | Pb (206) | 2s | Pb (208) | 2s | Th | 2s | U | 2s
3 | 10.1328 | 0.21494 | 65.6522 | 2.6802 | 1.1178 | 0.021729 | 9.156 | 0.19526 | 8.932 | 0.13673 | 41.2341 | 0.93683 | 63.4728 | 4.5148 | 43.7631 | 1.783 | 42.5059 | 0.7308 | 37.9093 | 2.1328 | 36.0617 | 0.61409 | 39.9582 | 1.6468 | 36.087 | 1.4421 | 84.2587 | 1.4904 | 42.6861 | 0.83698 | 41.985 | 0.86467 | 41.2273 | 0.7004 | 39.5875 | 0.81751 | 42.3323 | 0.86014 | 40.6653 | 0.65402 | 39.1003 | 1.1042 | 41.1719 | 1.096 | 39.0193 | 1.1234 | 42.9584 | 1.1558 | 42.3883 | 1.1098 | 40.1712 | 1.0092 | 38.8998 | 1.3355 | 40.9471 | 1.1478 | 42.4158 | 1.1777 | 39.6693 | 0.95028 | 41.9155 | 1.1927 | 39.3707 | 0.97751 | 39.0004 | 0.97548 | 40.0164 | 0.81564 | 37.2692 | 0.8675 | 39.7734 | 0.77012 | 40.7984 | 0.76114 | 40.1936 | 0.69268
4 | 10.058 | 0.21335 | 66.3107 | 2.6908 | 1.1079 | 0.021406 | 9.1013 | 0.18147 | 8.877 | 0.15959 | 39.6732 | 1.1842 | 70.6483 | 7.127 | 43.7904 | 1.4403 | 42.0398 | 0.77472 | 36.0528 | 2.0321 | 35.921 | 0.55981 | 39.5252 | 1.558 | 36.1938 | 1.83 | 83.757 | 1.5189 | 42.314 | 0.82567 | 42.4894 | 0.77831 | 41.3857 | 0.68417 | 39.392 | 0.80909 | 42.1686 | 0.80846 | 40.7032 | 0.563 | 39.3301 | 1.1579 | 41.1769 | 1.1165 | 38.8848 | 1.0407 | 41.6461 | 1.173 | 42.347 | 1.1969 | 39.9247 | 0.9247 | 38.443 | 1.3197 | 40.6597 | 1.0242 | 41.6877 | 1.1362 | 39.5194 | 0.93563 | 42.1933 | 1.2746 | 39.6711 | 1.0064 | 38.8776 | 0.85686 | 39.8425 | 0.71909 | 36.4022 | 0.77872 | 39.8027 | 0.75161 | 41.1876 | 0.71753 | 40.1257 | 0.69151
53 | 10.1296 | 0.21487 | 62.0748 | 2.4923 | 1.1283 | 0.02211 | 9.2172 | 0.23778 | 9.0207 | 0.16368 | 41.1935 | 0.86561 | 63.8924 | 6.4135 | 44.6146 | 1.9348 | 43.0378 | 0.73994 | 36.9754 | 1.7098 | 36.2702 | 0.55965 | 38.1351 | 1.7253 | 33.8306 | 1.7091 | 84.6822 | 1.6679 | 43.5749 | 1.0154 | 42.9256 | 0.99292 | 41.8585 | 0.77724 | 39.9661 | 0.97667 | 42.5007 | 0.9204 | 41.7943 | 0.82602 | 39.9091 | 1.0076 | 42.1339 | 1.308 | 39.4937 | 1.0536 | 43.6624 | 1.3263 | 43.3035 | 1.1186 | 40.803 | 0.98965 | 39.2336 | 1.2493 | 41.2965 | 1.1343 | 42.1797 | 1.2366 | 39.9912 | 0.95376 | 42.82 | 1.2 | 40.559 | 1.0567 | 39.5902 | 1.1223 | 40.7511 | 0.86725 | 36.8824 | 0.9193 | 40.1797 | 0.87353 | 41.8276 | 0.89045 | 40.4869 | 0.68887
54 | 10.1878 | 0.21611 | 66.9287 | 2.7076 | 1.1245 | 0.021727 | 9.2873 | 0.19867 | 9.0437 | 0.15584 | 40.4269 | 1.1081 | 70.4344 | 7.2641 | 44.6505 | 1.6828 | 42.9631 | 0.73866 | 37.2565 | 1.683 | 36.6592 | 0.56565 | 39.5139 | 1.8433 | 32.768 | 1.8129 | 85.1504 | 1.6113 | 43.2506 | 0.88654 | 42.7641 | 0.86509 | 42.1552 | 0.71627 | 39.9142 | 0.81072 | 42.2154 | 0.82815 | 41.8278 | 0.64664 | 39.8606 | 1.0986 | 42.1217 | 0.9721 | 39.6796 | 1.1467 | 43.313 | 1.4698 | 43.6132 | 1.3029 | 40.8563 | 0.94114 | 39.5002 | 1.2849 | 41.7632 | 1.1267 | 42.2569 | 1.2105 | 40.2918 | 1.0526 | 42.904 | 1.0684 | 40.6723 | 1.0215 | 39.6789 | 1.1462 | 40.9126 | 0.81712 | 36.9926 | 0.81846 | 40.2309 | 0.77856 | 41.8318 | 0.83576 | 40.6738 | 0.70096
103 | 10.1197 | 0.21466 | 62.8438 | 2.7519 | 1.115 | 0.02185 | 9.1318 | 0.1871 | 8.8846 | 0.13628 | 41.8452 | 1.0079 | 62.3116 | 4.5839 | 42.4478 | 1.4857 | 43.0482 | 0.77576 | 38.7253 | 1.7791 | 36.3419 | 0.61783 | 37.3777 | 1.8152 | 33.3844 | 1.645 | 83.5461 | 1.5267 | 42.7442 | 0.95286 | 41.918 | 0.82922 | 41.1814 | 0.68079 | 39.2955 | 0.73605 | 41.7111 | 0.82413 | 41.1087 | 0.59804 | 38.8492 | 1.0432 | 40.8417 | 1.4859 | 39.1372 | 1.0703 | 42.4124 | 1.1026 | 42.3679 | 1.0279 | 39.8739 | 0.91657 | 38.5915 | 1.2761 | 41.2041 | 1.0006 | 42.3708 | 1.2184 | 39.6505 | 1.0308 | 42.248 | 1.3297 | 39.8978 | 0.95667 | 38.5172 | 1.0293 | 40.1275 | 0.69796 | 37.0627 | 0.83958 | 40.0001 | 0.8849 | 41.0019 | 0.84019 | 39.721 | 0.68454
104 | 10.1623 | 0.21557 | 63.4454 | 3.2557 | 1.1196 | 0.021994 | 9.1481 | 0.21172 | 8.8908 | 0.13868 | 39.8319 | 0.89237 | 59.0117 | 4.3965 | 43.5323 | 1.7696 | 42.9097 | 0.73774 | 38.9021 | 1.9712 | 36.3218 | 0.56044 | 39.1455 | 1.6789 | 32.8525 | 1.5619 | 83.4454 | 1.4746 | 42.8431 | 0.89134 | 42.0177 | 0.85267 | 41.2837 | 0.68248 | 39.3872 | 0.79528 | 41.9888 | 0.81448 | 41.298 | 0.68862 | 39.2145 | 1.248 | 41.6211 | 1.0801 | 39.1107 | 1.1761 | 43.1619 | 1.3031 | 42.4347 | 1.1415 | 40.1858 | 0.90422 | 38.5043 | 1.5392 | 41.337 | 1.046 | 42.2537 | 1.0175 | 39.7417 | 0.91966 | 42.3975 | 1.1468 | 40.1158 | 1.0616 | 39.1577 | 1.0035 | 40.0143 | 0.73411 | 36.9889 | 0.83361 | 40.2374 | 0.96782 | 41.2437 | 0.9048 | 40.112 | 0.69128
153 | 10.2601 | 0.21764 | 62.4709 | 5.6024 | 1.1308 | 0.021848 | 9.2187 | 0.18381 | 8.9715 | 0.13512 | 41.8271 | 0.88122 | 58.2807 | 6.5932 | 44.6515 | 2.2081 | 42.6911 | 0.78135 | 41.3431 | 1.8438 | 36.705 | 0.56805 | 38.5867 | 1.3244 | 33.8476 | 1.631 | 85.0003 | 1.4933 | 43.4421 | 0.93528 | 42.5588 | 0.89122 | 41.8537 | 0.75582 | 39.5925 | 0.91934 | 42.229 | 0.98781 | 41.4552 | 0.65346 | 39.3131 | 0.96529 | 41.497 | 1.2064 | 39.2637 | 1.1387 | 43.1638 | 1.483 | 43.4569 | 1.1491 | 40.8666 | 1.0207 | 38.8082 | 1.3436 | 41.2993 | 1.0334 | 41.973 | 1.2072 | 40.3718 | 1.0184 | 42.4946 | 1.4013 | 40.0067 | 1.0931 | 39.0269 | 1.0744 | 40.3117 | 0.76518 | 37.2117 | 1.0308 | 40.3253 | 0.9099 | 41.379 | 0.85084 | 40.1269 | 0.7197
154 | 10.2099 | 0.21658 | 61.2537 | 5.5644 | 1.1338 | 0.021906 | 9.175 | 0.21507 | 8.9453 | 0.14107 | 39.388 | 1.1267 | 62.1491 | 5.2503 | 43.7559 | 1.9254 | 42.5753 | 0.80157 | 41.4028 | 2.0519 | 36.8062 | 0.56792 | 38.542 | 1.5876 | 33.6413 | 1.2295 | 85.0197 | 1.5081 | 43.6989 | 0.9083 | 43.0634 | 0.92376 | 42.1198 | 0.71782 | 39.8244 | 0.97924 | 42.4426 | 0.86888 | 41.4919 | 0.73611 | 39.3288 | 1.1925 | 41.4342 | 1.1109 | 39.502 | 1.1127 | 43.4718 | 1.2573 | 43.8203 | 1.2857 | 41.1844 | 1.0245 | 39.2095 | 1.3712 | 41.5838 | 1.1666 | 42.8728 | 1.2232 | 40.5125 | 1.0168 | 42.7505 | 1.183 | 40.5629 | 1.0899 | 39.354 | 1.1845 | 40.5319 | 0.71621 | 37.2226 | 1.0237 | 40.712 | 0.84087 | 41.6333 | 0.86257 | 40.4421 | 0.69696
203 | 10.295 | 0.21838 | 57.1601 | 5.2237 | 1.1329 | 0.021889 | 9.2444 | 0.21316 | 9.0727 | 0.13192 | 39.995 | 1.2547 | 58.7507 | 4.3531 | 44.0391 | 2.0289 | 42.4043 | 0.85003 | 38.7277 | 1.6086 | 36.7485 | 0.57925 | 39.1192 | 1.9224 | 33.7962 | 1.6144 | 85.8109 | 1.5208 | 43.4624 | 0.80094 | 43.1964 | 0.95538 | 41.6845 | 0.68911 | 39.4155 | 0.87803 | 42.162 | 0.80891 | 41.2285 | 0.74794 | 39.2677 | 1.0443 | 41.9682 | 1.1791 | 39.4425 | 0.99518 | 42.9238 | 1.2923 | 43.9194 | 1.0085 | 41.0044 | 1.0159 | 39.5647 | 1.3477 | 41.6363 | 1.0232 | 42.5027 | 1.1702 | 40.5251 | 0.97279 | 42.8507 | 1.067 | 40.5577 | 1.0166 | 39.9165 | 1.056 | 40.3295 | 0.86777 | 37.5017 | 0.93691 | 40.8217 | 0.89664 | 41.444 | 0.75475 | 40.3913 | 0.68724
204 | 10.1674 | 0.21567 | 61.1829 | 5.5744 | 1.1317 | 0.022005 | 9.1841 | 0.18678 | 9.0507 | 0.17735 | 39.6425 | 1.0887 | 61.9447 | 6.6938 | 42.904 | 1.7595 | 42.6871 | 0.73704 | 38.8454 | 1.8322 | 36.6675 | 0.63553 | 39.4926 | 1.4986 | 34.0954 | 1.4841 | 85.2576 | 1.7404 | 43.3337 | 0.81578 | 42.9798 | 1.0582 | 41.5152 | 0.8399 | 39.5181 | 0.98274 | 42.2428 | 0.87569 | 41.2076 | 0.75572 | 39.1632 | 1.1279 | 40.8349 | 1.2894 | 38.9713 | 1.1474 | 43.0676 | 0.99679 | 42.9124 | 1.134 | 40.8588 | 0.96963 | 39.5686 | 1.2145 | 41.3995 | 1.0845 | 42.0327 | 1.1085 | 40.1098 | 1.0109 | 42.7432 | 1.1636 | 40.5635 | 1.0915 | 39.2384 | 1.0685 | 40.1895 | 0.75383 | 36.5361 | 0.94236 | 40.6727 | 0.8853 | 41.261 | 0.82087 | 40.3579 | 0.69551
226 | 10.2356 | 0.21712 | 66.1225 | 6.25 | 1.1361 | 0.022082 | 9.2038 | 0.19663 | 8.9454 | 0.12507 | 38.7281 | 1.155 | 57.6807 | 8.7038 | 42.6245 | 1.8552 | 42.7993 | 0.73973 | 36.8295 | 1.6252 | 36.5786 | 0.5644 | 38.1057 | 1.8089 | 34.6874 | 1.6126 | 84.2948 | 1.5516 | 43.3293 | 0.79103 | 42.4598 | 0.90852 | 41.5653 | 0.68714 | 39.706 | 0.78769 | 42.0612 | 0.76929 | 41.4934 | 0.62074 | 39.0141 | 1.2057 | 41.7822 | 1.1757 | 39.344 | 0.97871 | 43.415 | 1.1543 | 43.0505 | 1.0995 | 40.016 | 0.94515 | 38.8649 | 1.4034 | 41.8225 | 1.0455 | 43.0023 | 1.0953 | 40.26 | 0.98789 | 42.8072 | 1.0529 | 40.62 | 0.97663 | 39.337 | 1.0163 | 40.3242 | 0.81673 | 37.5544 | 1.0011 | 40.3784 | 0.82526 | 41.5047 | 0.88495 | 40.5791 | 0.70475
227 | 10.2785 | 0.21803 | 59.3874 | 4.8489 | 1.1358 | 0.021945 | 9.1704 | 0.1993 | 8.9689 | 0.18088 | 40.9197 | 1.189 | 56.1356 | 6.6351 | 44.1517 | 2.4045 | 42.8866 | 0.74938 | 36.8567 | 1.6883 | 36.5449 | 0.5669 | 37.1276 | 1.7004 | 32.8108 | 1.4818 | 84.4347 | 1.6648 | 43.1808 | 0.94525 | 42.2753 | 0.90271 | 42.1237 | 0.70421 | 39.8877 | 0.8514 | 42.3619 | 0.89151 | 41.5326 | 0.67452 | 39.709 | 1.2055 | 42.7666 | 1.1785 | 39.1842 | 1.0622 | 43.2492 | 1.0438 | 42.4001 | 1.1275 | 39.8784 | 0.9176 | 38.7337 | 1.1909 | 41.4993 | 1.1374 | 43.1356 | 1.0206 | 40.2702 | 0.90258 | 42.9607 | 1.2965 | 40.6144 | 1.0073 | 39.7239 | 0.93058 | 40.7699 | 0.75694 | 37.3306 | 0.87498 | 40.3427 | 0.79764 | 41.3092 | 0.70618 | 40.5166 | 0.69825
Average | 10.1864 |  | 62.9028 |  | 1.1262 |  | 9.1865 |  | 8.9669 |  | 40.3921 |  | 62.0594 |  | 43.7438 |  | 42.7124 |  | 38.3189 |  | 36.4689 |  | 38.7191 |  | 33.9996 |  | 84.5548 |  | 43.155 |  | 42.5528 |  | 41.6628 |  | 39.6239 |  | 42.2014 |  | 41.3172 |  | 39.3383 |  | 41.6125 |  | 39.2528 |  | 43.0371 |  | 43.0012 |  | 40.4686 |  | 38.9935 |  | 41.3707 |  | 42.3903 |  | 40.0761 |  | 42.5904 |  | 40.2677 |  | 39.2849 |  | 40.3434 |  | 37.0796 |  | 40.2897 |  | 41.3685 |  | 40.3106
2s | 0.14356 |  | 5.9675 |  | 0.018326 |  | 0.10243 |  | 0.13462 |  | 2.0117 |  | 9.3056 |  | 1.5188 |  | 0.59332 |  | 3.4218 |  | 0.56711 |  | 1.7945 |  | 2.2963 |  | 1.4662 |  | 0.83437 |  | 0.88263 |  | 0.7166 |  | 0.46419 |  | 0.42691 |  | 0.73384 |  | 0.65747 |  | 1.1514 |  | 0.48773 |  | 1.0822 |  | 1.2188 |  | 0.99878 |  | 0.82 |  | 0.66492 |  | 0.86404 |  | 0.7069 |  | 0.66843 |  | 0.88219 |  | 0.80197 |  | 0.67303 |  | 0.70059 |  | 0.66763 |  | 0.61577 |  | 0.52992
GeoRem Preferred value | 10.2 |  | 68 |  | 1.15 |  | 8.5 |  | 8.5 |  | 38.7 |  | 51 |  | 44 |  | 38.8 |  | 36.4 |  | 35.5 |  | 38.8 |  | 39.1 |  | 78.4 |  | 38.3 |  | 37.9 |  | 38.9 |  | 36 |  | 38.4 |  | 37.9 |  | 35.5 |  | 37.7 |  | 35.6 |  | 37.3 |  | 37.3 |  | 37.6 |  | 35.5 |  | 38.3 |  | 38 |  | 36.8 |  | 39.2 |  | 37 |  | 36.7 |  | 37.6 |  | 38.6 |  | 38.6 |  | 37.79 |  | 37.38
2s | 0.2 |  | 5.1 |  | 0.02 |  | 0.07 |  | 0.07 |  | 0.9 |  | 2 |  | 2.3 |  | 1.2 |  | 1.5 |  | 1 |  | 0.2 |  | 1.7 |  | 0.2 |  | 1.4 |  | 1.2 |  | 2.1 |  | 0.7 |  | 0.7 |  | 1 |  | 0.7 |  | 0.8 |  | 0.8 |  | 0.9 |  | 0.9 |  | 1.1 |  | 0.7 |  | 0.8 |  | 0.9 |  | 0.6 |  | 0.9 |  | 0.9 |  | 1.2 |  | 1.9 |  | 0.2 |  | 0.2 |  | 0.08 |  | 0.08

## BCR-1
Table 4f. LA-ICP-MS trace element analyses (33 μm) of secondary reference glass BCR-1.
| Major elements (wt %) | [13 empty] | Trace elements (μg/g)
Analyses | Na | 2s | Mg | 2s | Al | 2s | Ca (43) | 2s | Ca (44) | 2s | Mn | 2s | Fe | 2s | Ti | 2s | V | 2s | Cr | 2s | Co | 2s | Ni | 2s | Zn | 2s | Sr | 2s | Y | 2s | Zr | 2s | Nb | 2s | La | 2s | Ce | 2s | Pr | 2s | Nd | 2s | Sm | 2s | Eu | 2s | Gd (157) | 2s | Gd (158) | 2s | Tb | 2s | Dy | 2s | Ho | 2s | Er | 2s | Tm | 2s | Yb | 2s | Lu | 2s | Hf | 2s | Ta | 2s | Pb (206) | 2s | Pb (208) | 2s | Th | 2s | U | 2s
5 | 2.2134 | 0.062678 | 2.0824 | 0.053012 | 6.8178 | 0.18183 | 5.14 | 0.12733 | 5.0824 | 0.11126 | 0.15251 | 0.0039247 | 9.9014 | 0.22202 | 13157.4099 | 501.425 | 434.366 | 10.6882 | 14.9841 | 1.2827 | 38.6341 | 1.2684 | 11.6884 | 1.4412 | 158.7486 | 5.3611 | 334.5889 | 9.0031 | 33.7078 | 0.59976 | 178.1319 | 4.4471 | 11.5499 | 0.33452 | 24.5594 | 0.56739 | 52.5864 | 1.7462 | 6.4342 | 0.18734 | 27.8154 | 0.69654 | 6.479 | 0.28196 | 1.9042 | 0.081397 | 6.588 | 0.29185 | 6.7545 | 0.24332 | 0.97597 | 0.043269 | 6.1564 | 0.25956 | 1.239 | 0.04021 | 3.4461 | 0.1808 | 0.48713 | 0.032461 | 3.286 | 0.14807 | 0.48836 | 0.032339 | 4.8115 | 0.1806 | 0.68017 | 0.031749 | 10.3424 | 0.29772 | 11.2101 | 0.3789 | 5.8313 | 0.1492 | 1.6624 | 0.051073
6 | 2.2554 | 0.047843 | 2.1145 | 0.05101 | 6.9527 | 0.16061 | 5.2534 | 0.11907 | 5.1386 | 0.1053 | 0.15421 | 0.0031535 | 10.0606 | 0.15503 | 13194.7588 | 283.9093 | 435.8643 | 7.9083 | 13.8939 | 1.8798 | 38.2964 | 0.59091 | 13.5933 | 1.4183 | 161.142 | 6.2408 | 336.2139 | 6.0957 | 34.0615 | 0.65631 | 182.7413 | 3.1763 | 11.7856 | 0.28383 | 25.0193 | 0.49654 | 52.6767 | 1.1444 | 6.5437 | 0.15169 | 28.5277 | 0.89138 | 6.4712 | 0.24479 | 1.9315 | 0.092127 | 6.7425 | 0.26445 | 6.824 | 0.29977 | 0.98881 | 0.04078 | 6.1905 | 0.31426 | 1.2774 | 0.051178 | 3.7085 | 0.17159 | 0.51283 | 0.022599 | 3.2199 | 0.15565 | 0.50993 | 0.023973 | 4.8769 | 0.17331 | 0.72463 | 0.03243 | 10.6185 | 0.37156 | 10.994 | 0.32299 | 5.8916 | 0.16013 | 1.6992 | 0.042221
55 | 2.304 | 0.06418 | 2.1494 | 0.051227 | 7.0117 | 0.14013 | 5.2432 | 0.11621 | 5.2452 | 0.085349 | 0.15773 | 0.0033463 | 10.0845 | 0.14512 | 13291.9391 | 242.1634 | 444.589 | 9.3739 | 16.952 | 1.3637 | 38.1692 | 0.58895 | 10.8464 | 1.5907 | 157.8222 | 5.4134 | 339.1323 | 6.6697 | 35.2432 | 0.73 | 182.7972 | 3.3769 | 11.861 | 0.25665 | 25.4986 | 0.62269 | 53.0882 | 0.97097 | 6.724 | 0.13296 | 29.2335 | 0.80851 | 6.6506 | 0.24019 | 1.997 | 0.10551 | 6.9175 | 0.29278 | 6.9799 | 0.27377 | 0.98673 | 0.042148 | 6.4587 | 0.28593 | 1.2334 | 0.041924 | 3.8829 | 0.16057 | 0.5011 | 0.025715 | 3.4331 | 0.22124 | 0.51298 | 0.032781 | 4.9306 | 0.21222 | 0.71078 | 0.031514 | 10.8735 | 0.32987 | 11.2255 | 0.30388 | 5.9095 | 0.13223 | 1.7076 | 0.050875
56 | 2.258 | 0.055744 | 2.162 | 0.039489 | 7.0591 | 0.1801 | 5.4599 | 0.48865 | 5.307 | 0.10238 | 0.15862 | 0.0042278 | 10.0219 | 0.17928 | 13317.9201 | 306.2393 | 446.1227 | 11.7476 | 14.5665 | 1.4112 | 38.0379 | 0.60722 | 12.8866 | 2.9697 | 173.0711 | 16.3759 | 341.0723 | 6.9359 | 35.1289 | 0.81608 | 183.8096 | 4.1208 | 11.8837 | 0.33812 | 25.504 | 0.67444 | 52.6806 | 1.3269 | 6.9005 | 0.2055 | 29.5483 | 1.0152 | 6.7351 | 0.30539 | 2.0623 | 0.1044 | 6.8341 | 0.36168 | 6.9978 | 0.2752 | 1.006 | 0.038321 | 6.5867 | 0.32166 | 1.2827 | 0.042022 | 3.724 | 0.18411 | 0.51922 | 0.035821 | 3.4609 | 0.16209 | 0.50458 | 0.023975 | 4.8917 | 0.20557 | 0.71196 | 0.039222 | 10.9523 | 0.32853 | 11.5188 | 0.38971 | 6.0426 | 0.15766 | 1.7063 | 0.055804
105 | 2.2832 | 0.059294 | 2.1451 | 0.053286 | 7.0521 | 0.17642 | 5.2425 | 0.14055 | 5.1819 | 0.095908 | 0.15635 | 0.003357 | 9.9048 | 0.1786 | 13106.4554 | 293.1852 | 442.283 | 9.634 | 17.5064 | 1.2618 | 38.1389 | 0.71013 | 11.6472 | 1.179 | 154.9467 | 5.3658 | 333.3527 | 6.9714 | 34.4378 | 0.79209 | 178.396 | 3.7003 | 11.6629 | 0.29633 | 25.0181 | 0.49059 | 52.0436 | 1.0676 | 6.6775 | 0.17514 | 28.8749 | 1.0187 | 6.5903 | 0.25344 | 1.9819 | 0.09165 | 7.0631 | 0.3083 | 6.8595 | 0.28896 | 0.99489 | 0.056341 | 6.1065 | 0.24845 | 1.2498 | 0.049235 | 3.6969 | 0.16132 | 0.49778 | 0.030387 | 3.3347 | 0.18405 | 0.50994 | 0.039756 | 4.6865 | 0.1524 | 0.69678 | 0.032927 | 10.902 | 0.4582 | 11.2614 | 0.2448 | 5.7969 | 0.15795 | 1.718 | 0.044852
106 | 2.2508 | 0.052679 | 2.0924 | 0.05033 | 6.9154 | 0.14157 | 5.1824 | 0.11409 | 5.1305 | 0.098254 | 0.15483 | 0.0028767 | 9.8397 | 0.1416 | 12979.5786 | 242.3521 | 438.2148 | 7.5341 | 16.9659 | 1.2995 | 37.9077 | 0.63764 | 12.6859 | 1.0175 | 151.5543 | 5.0172 | 333.9465 | 6.4364 | 34.5455 | 0.66616 | 178.2209 | 2.9771 | 11.7128 | 0.21905 | 24.9355 | 0.46918 | 51.7525 | 1.0604 | 6.5122 | 0.19276 | 28.178 | 0.82758 | 6.5142 | 0.33816 | 1.9798 | 0.077549 | 6.9951 | 0.34183 | 6.7605 | 0.30868 | 0.99335 | 0.038677 | 6.2637 | 0.2852 | 1.2588 | 0.049548 | 3.7728 | 0.16868 | 0.526 | 0.036609 | 3.3785 | 0.18339 | 0.50028 | 0.032749 | 4.7145 | 0.2341 | 0.6808 | 0.026923 | 10.7363 | 0.39105 | 11.2692 | 0.31434 | 5.878 | 0.13273 | 1.7075 | 0.043902
155 | 2.278 | 0.052728 | 2.0698 | 0.04817 | 7.1216 | 0.21069 | 5.2413 | 0.12837 | 5.1612 | 0.096182 | 0.1589 | 0.0032011 | 10.0009 | 0.14392 | 13251.6105 | 266.8319 | 443.3726 | 7.7424 | 19.2152 | 1.3047 | 38.5599 | 0.76113 | 11.7707 | 1.551 | 155.8791 | 6.1035 | 340.9581 | 7.4568 | 34.9469 | 0.80252 | 184.0009 | 3.5687 | 11.7921 | 0.31202 | 26.9057 | 3.4575 | 53.2066 | 1.2744 | 6.7145 | 0.16486 | 28.6385 | 0.95432 | 6.6058 | 0.26079 | 1.9562 | 0.081444 | 6.819 | 0.31744 | 6.9213 | 0.24603 | 1.0251 | 0.048575 | 6.3737 | 0.30069 | 1.2637 | 0.048371 | 3.7652 | 0.14599 | 0.4989 | 0.026774 | 3.5534 | 0.17323 | 0.48142 | 0.027158 | 4.8021 | 0.18057 | 0.73304 | 0.02916 | 10.9874 | 0.5286 | 11.1697 | 0.30713 | 5.8969 | 0.17019 | 1.7865 | 0.067914
156 | 2.2743 | 0.14065 | 2.1141 | 0.07295 | 6.9163 | 0.32089 | 5.1241 | 0.18227 | 5.0861 | 0.21955 | 0.15558 | 0.0062618 | 9.9062 | 0.41868 | 13063.8175 | 626.4958 | 437.106 | 8.4983 | 17.8524 | 1.2886 | 37.577 | 1.3506 | 11.093 | 1.5045 | 154.7241 | 5.947 | 333.2613 | 13.82 | 34.1943 | 1.3725 | 182.9921 | 7.2242 | 11.5836 | 0.42424 | 24.8832 | 0.99263 | 52.1358 | 2.0893 | 6.6465 | 0.21337 | 28.6437 | 1.4239 | 6.5642 | 0.34832 | 1.9871 | 0.11275 | 6.7729 | 0.35506 | 7.1052 | 0.25106 | 1.0028 | 0.046393 | 6.4828 | 0.34674 | 1.239 | 0.084527 | 3.8069 | 0.17996 | 0.51484 | 0.033896 | 3.3728 | 0.22011 | 0.49896 | 0.035275 | 4.7379 | 0.24657 | 0.70638 | 0.038487 | 10.8403 | 0.50749 | 11.148 | 0.47201 | 5.8755 | 0.25458 | 1.6801 | 0.086387
205 | 2.2647 | 0.056891 | 2.1207 | 0.043001 | 6.9912 | 0.1606 | 5.2374 | 0.14242 | 5.1538 | 0.11625 | 0.15743 | 0.0034653 | 9.8623 | 0.18461 | 13302.0996 | 360.1128 | 432.1075 | 8.0739 | 18.1776 | 1.548 | 38.1776 | 0.69186 | 11.8272 | 1.862 | 155.5036 | 5.7745 | 339.2224 | 7.7529 | 34.6172 | 0.71698 | 181.8114 | 3.5809 | 11.6759 | 0.29833 | 24.9811 | 0.58187 | 52.4038 | 1.0527 | 6.5374 | 0.14432 | 28.176 | 0.7858 | 6.5805 | 0.33902 | 1.9939 | 0.10272 | 6.834 | 0.34369 | 6.9025 | 0.25232 | 1.0279 | 0.03771 | 6.3224 | 0.28851 | 1.2268 | 0.058183 | 3.7633 | 0.15056 | 0.49754 | 0.030461 | 3.3092 | 0.13632 | 0.50948 | 0.027737 | 4.7072 | 0.16818 | 0.70854 | 0.02789 | 10.5999 | 0.34685 | 10.9942 | 0.30894 | 5.8045 | 0.15 | 1.6862 | 0.061706
206 | 2.3115 | 0.075692 | 2.1518 | 0.064017 | 7.0995 | 0.19377 | 5.1892 | 0.14107 | 5.1907 | 0.11322 | 0.15835 | 0.0036542 | 9.9553 | 0.17393 | 13432.2505 | 340.8194 | 438.6035 | 9.725 | 18.0551 | 1.1091 | 38.2652 | 0.83962 | 13.2241 | 1.1359 | 152.5615 | 6.0462 | 341.0539 | 7.5988 | 35.0685 | 0.80685 | 184.727 | 4.2514 | 11.6487 | 0.33852 | 25.1144 | 0.62351 | 52.8045 | 1.2883 | 6.6306 | 0.15692 | 28.2126 | 1.0404 | 6.9899 | 0.302 | 2.0205 | 0.11474 | 6.8602 | 0.3169 | 7.0576 | 0.25466 | 1.0093 | 0.036731 | 6.4627 | 0.26826 | 1.3063 | 0.052641 | 3.8771 | 0.23845 | 0.55308 | 0.027097 | 3.4844 | 0.1821 | 0.47098 | 0.033059 | 4.8626 | 0.23935 | 0.67826 | 0.032631 | 10.6049 | 0.34941 | 11.2693 | 0.36649 | 6.0023 | 0.21701 | 1.6411 | 0.047247
228 | 2.2925 | 0.051037 | 2.0953 | 0.044307 | 7.1221 | 0.15286 | 5.2076 | 0.12451 | 5.1957 | 0.15222 | 0.15479 | 0.0032499 | 9.8986 | 0.14244 | 13063.4357 | 296.7836 | 439.1915 | 7.643 | 17.3446 | 1.1546 | 38.4177 | 0.66763 | 12.9461 | 1.4956 | 151.2129 | 5.0247 | 335.7765 | 7.2773 | 34.6487 | 0.6619 | 181.8253 | 3.5201 | 11.8541 | 0.25661 | 25.1127 | 0.52357 | 52.782 | 1.1645 | 6.6139 | 0.1608 | 28.8583 | 0.84019 | 6.786 | 0.27797 | 1.993 | 0.098317 | 7.1017 | 0.34945 | 6.835 | 0.25772 | 1.0394 | 0.041446 | 6.38 | 0.29076 | 1.3004 | 0.071403 | 3.7578 | 0.17255 | 0.49757 | 0.028957 | 3.4112 | 0.16141 | 0.49935 | 0.026251 | 4.7071 | 0.18945 | 0.70837 | 0.027648 | 11.1813 | 0.31422 | 11.2895 | 0.24461 | 5.9007 | 0.14493 | 1.729 | 0.068816
229 | 2.2503 | 0.053174 | 2.1489 | 0.065116 | 6.9892 | 0.14954 | 5.1365 | 0.1324 | 5.1314 | 0.11409 | 0.1532 | 0.0036034 | 9.9028 | 0.18001 | 12853.9433 | 272.7893 | 442.6273 | 9.4055 | 16.8458 | 1.0027 | 37.6111 | 0.84691 | 11.064 | 1.3021 | 154.8367 | 6.1153 | 332.5357 | 6.3613 | 34.6608 | 0.69413 | 180.619 | 3.5682 | 12.0931 | 0.34379 | 25.0511 | 0.56681 | 52.434 | 1.0733 | 6.6088 | 0.16857 | 28.536 | 0.86244 | 6.8349 | 0.27797 | 1.9365 | 0.098728 | 6.7241 | 0.30351 | 6.6937 | 0.27546 | 0.97056 | 0.032983 | 6.3159 | 0.31997 | 1.2847 | 0.048115 | 3.6125 | 0.18116 | 0.50043 | 0.035747 | 3.4159 | 0.15434 | 0.49317 | 0.032938 | 4.8022 | 0.16244 | 0.71285 | 0.025061 | 11.0589 | 0.415 | 11.3183 | 0.26696 | 5.7619 | 0.13525 | 1.7006 | 0.057071
Average | 2.2697 |  | 2.1206 |  | 7.0041 |  | 5.2215 |  | 5.1671 |  | 0.15604 |  | 9.9449 |  | 13167.9349 |  | 439.5373 |  | 16.8633 |  | 38.1494 |  | 12.1061 |  | 156.8336 |  | 336.7595 |  | 34.6051 |  | 181.6727 |  | 11.7586 |  | 25.2153 |  | 52.5496 |  | 6.6287 |  | 28.6036 |  | 6.6502 |  | 1.9786 |  | 6.8544 |  | 6.891 |  | 1.0017 |  | 6.3417 |  | 1.2635 |  | 3.7345 |  | 0.50887 |  | 3.3883 |  | 0.49829 |  | 4.7942 |  | 0.70438 |  | 10.8081 |  | 11.2223 |  | 5.8826 |  | 1.702
2s | 0.05388 |  | 0.061751 |  | 0.18637 |  | 0.17636 |  | 0.12743 |  | 0.0043542 |  | 0.15824 |  | 326.5113 |  | 8.6291 |  | 3.1878 |  | 0.66056 |  | 1.8476 |  | 11.7435 |  | 6.6589 |  | 0.91836 |  | 4.6634 |  | 0.30352 |  | 1.1785 |  | 0.84695 |  | 0.24167 |  | 0.96841 |  | 0.31498 |  | 0.084567 |  | 0.29399 |  | 0.25382 |  | 0.042079 |  | 0.28943 |  | 0.053526 |  | 0.23482 |  | 0.03562 |  | 0.18422 |  | 0.025675 |  | 0.16661 |  | 0.03475 |  | 0.47171 |  | 0.28296 |  | 0.16176 |  | 0.07191
GeoRem Preferred value | 2.47 |  | 2.09 |  | 7.71 |  | 4.99 |  | 4.99 |  | 0.14 |  | 9.39 |  | 13437.2341 |  | 404.4 |  | 13.5 |  | 37.55 |  | 11.68 |  | 128.5 |  | 334.9 |  | 36.91 |  | 190.3 |  | 12.74 |  | 25.46 |  | 53.94 |  | 6.765 |  | 28.68 |  | 6.603 |  | 1.957 |  | 6.725 |  | 6.725 |  | 1.063 |  | 6.391 |  | 1.268 |  | 3.658 |  | 0.535 |  | 3.377 |  | 0.4988 |  | 4.923 |  | 0.786 |  | 13.44 |  | 13.44 |  | 5.79 |  | 1.678
2s | 0.04 |  | 0.02 |  | 0.05 |  | 0.06 |  | 0.06 |  | 0.002 |  | 0.1 |  | 209.7695 |  | 5.2 |  | 1.3 |  | 0.64 |  | 0.84 |  | 2.9 |  | 3.5 |  | 0.55 |  | 2.2 |  | 0.23 |  | 0.65 |  | 0.34 |  | 0.066 |  | 0.1 |  | 0.022 |  | 0.018 |  | 0.054 |  | 0.054 |  | 0.013 |  | 0.075 |  | 0.016 |  | 0.056 |  | 0.012 |  | 0.037 |  | 0.0048 |  | 0.06 |  | 0.021 |  | 0.23 |  | 0.23 |  | 0.11 |  | 0.019

## BIR-1G
Table 4g. LA-ICP-MS trace element analyses (33 μm) of secondary reference glass BIR-1G.
| Major elements (wt %) | [13 empty] | Trace elements (μg/g)
Analyses | Na | 2s | Mg | 2s | Al | 2s | Ca (43) | 2s | Ca (44) | 2s | Mn | 2s | Fe | 2s | Ti | 2s | V | 2s | Cr | 2s | Co | 2s | Ni | 2s | Zn | 2s | Sr | 2s | Y | 2s | Zr | 2s | Nb | 2s | La | 2s | Ce | 2s | Pr | 2s | Nd | 2s | Sm | 2s | Eu | 2s | Gd (157) | 2s | Gd (158) | 2s | Tb | 2s | Dy | 2s | Ho | 2s | Er | 2s | Tm | 2s | Yb | 2s | Lu | 2s | Hf | 2s | Ta | 2s | Pb (206) | 2s | Pb (208) | 2s | Th | 2s | U | 2s
7 | 1.3469 | 0.050312 | 5.4336 | 0.12848 | 7.6663 | 0.23232 | 8.9831 | 0.23621 | 8.8603 | 0.21851 | 0.13247 | 0.002803 | 8.2317 | 0.12673 | 5232.3538 | 127.2625 | 320.1747 | 7.272 | 350.1813 | 14.9862 | 54.3144 | 0.98225 | 179.3372 | 4.9461 | 83.3537 | 3.8171 | 100.8453 | 2.296 | 13.7832 | 0.32381 | 13.1165 | 0.38139 | 0.4556 | 0.027435 | 0.53975 | 0.038491 | 1.817 | 0.079533 | 0.35496 | 0.026045 | 2.2083 | 0.13215 | 1.0023 | 0.10921 | 0.47851 | 0.040288 | 1.7527 | 0.15352 | 1.8478 | 0.12054 | 0.30878 | 0.027269 | 2.2852 | 0.14799 | 0.51401 | 0.027038 | 1.4965 | 0.082578 | 0.21967 | 0.022427 | 1.5135 | 0.13724 | 0.23836 | 0.019707 | 0.54811 | 0.058024 | 0.031076 | 0.0049533 | 3.7075 | 0.21584 | 3.7279 | 0.16939 | 0.023483 | 0.0050079 | 0.019282 | 0.0053108
8 | 1.3277 | 0.035844 | 5.4262 | 0.13194 | 7.5969 | 0.20966 | 8.9676 | 0.27973 | 8.8069 | 0.22142 | 0.13225 | 0.0030041 | 8.2385 | 0.14896 | 5247.5357 | 133.3376 | 320.0801 | 7.0209 | 348.7463 | 14.3481 | 54.0857 | 0.89804 | 179.1875 | 5.2761 | 83.5035 | 3.2246 | 102.4862 | 2.4964 | 13.9966 | 0.41888 | 13.0598 | 0.51052 | 0.48275 | 0.031562 | 0.53398 | 0.029017 | 1.776 | 0.063818 | 0.33332 | 0.029824 | 2.1757 | 0.1422 | 0.97869 | 0.10436 | 0.48926 | 0.039398 | 1.8337 | 0.14784 | 1.9052 | 0.14107 | 0.32623 | 0.02441 | 2.3974 | 0.16932 | 0.51486 | 0.031986 | 1.557 | 0.12057 | 0.21426 | 0.017891 | 1.5971 | 0.11374 | 0.21308 | 0.018736 | 0.52825 | 0.049692 | 0.029781 | 0.0053779 | 3.7756 | 0.14117 | 3.7061 | 0.13859 | 0.029403 | 0.0070044 | 0.01917 | 0.0064282
57 | 1.3656 | 0.060631 | 5.6225 | 0.1478 | 7.772 | 0.21316 | 9.1264 | 0.22401 | 8.9973 | 0.18117 | 0.13662 | 0.0031988 | 8.3067 | 0.13539 | 5325.09 | 127.7736 | 328.1496 | 8.0036 | 338.1784 | 13.438 | 54.3053 | 0.93757 | 188.3014 | 5.369 | 84.3317 | 3.7708 | 102.6369 | 2.126 | 14.2637 | 0.3708 | 13.3451 | 0.50087 | 0.45969 | 0.028484 | 0.56585 | 0.037551 | 1.797 | 0.054586 | 0.36388 | 0.024869 | 2.288 | 0.17001 | 1.1198 | 0.094535 | 0.49983 | 0.041251 | 1.7134 | 0.11104 | 1.7709 | 0.10905 | 0.33029 | 0.018143 | 2.3798 | 0.2344 | 0.52045 | 0.034165 | 1.6226 | 0.091737 | 0.22199 | 0.01526 | 1.7252 | 0.13564 | 0.21942 | 0.017631 | 0.52358 | 0.053485 | 0.033908 | 0.008264 | 3.769 | 0.17826 | 3.871 | 0.17646 | 0.025254 | 0.0067811 | 0.020638 | 0.0037087
58 | 1.3292 | 0.031873 | 5.5478 | 0.11408 | 7.8012 | 0.21273 | 9.0042 | 0.22238 | 8.8422 | 0.1751 | 0.13673 | 0.0029107 | 8.3764 | 0.15447 | 5333.2996 | 142.1741 | 326.2617 | 5.9814 | 339.653 | 13.163 | 54.4612 | 0.89481 | 188.3908 | 5.3593 | 82.0246 | 3.3825 | 101.7618 | 2.3917 | 14.1822 | 0.36784 | 13.0479 | 0.30804 | 0.44614 | 0.032115 | 0.53542 | 0.036018 | 1.7891 | 0.061466 | 0.33177 | 0.027557 | 2.1698 | 0.17458 | 1.0363 | 0.14987 | 0.47015 | 0.050059 | 1.8012 | 0.11309 | 1.8698 | 0.14388 | 0.33418 | 0.025003 | 2.4014 | 0.12651 | 0.54503 | 0.032202 | 1.5749 | 0.12413 | 0.23129 | 0.026453 | 1.5135 | 0.14087 | 0.23118 | 0.018214 | 0.53285 | 0.057017 | 0.027567 | 0.007584 | 3.7432 | 0.17798 | 3.8083 | 0.14035 | 0.025177 | 0.0055787 | 0.017921 | 0.0035671
107 | 1.3486 | 0.033646 | 5.5333 | 0.099854 | 7.6464 | 0.19116 | 9.077 | 0.23066 | 8.8383 | 0.23304 | 0.13551 | 0.0025177 | 8.2838 | 0.22078 | 5237.1418 | 126.1654 | 325.1381 | 6.1602 | 349.8214 | 14.0103 | 53.8727 | 1.044 | 182.1964 | 5.4471 | 81.9569 | 4.3027 | 100.4832 | 2.6534 | 14.2397 | 0.40581 | 13.1538 | 0.36002 | 0.46459 | 0.039128 | 0.56232 | 0.027219 | 1.7955 | 0.064762 | 0.34548 | 0.027018 | 2.2168 | 0.18784 | 0.94337 | 0.09041 | 0.5103 | 0.050223 | 1.7413 | 0.14246 | 1.8034 | 0.11592 | 0.3214 | 0.02084 | 2.468 | 0.15258 | 0.51476 | 0.032096 | 1.5135 | 0.10561 | 0.22589 | 0.015411 | 1.5294 | 0.11107 | 0.22144 | 0.015547 | 0.457 | 0.043671 | 0.033864 | 0.0050893 | 3.8279 | 0.21324 | 3.6232 | 0.1127 | 0.029353 | 0.0060615 | 0.016767 | 0.0034059
108 | 1.3495 | 0.037692 | 5.5098 | 0.11914 | 7.702 | 0.21732 | 9.0266 | 0.19909 | 8.6862 | 0.21438 | 0.13486 | 0.0028683 | 8.2344 | 0.15965 | 5231.229 | 133.1184 | 327.3784 | 5.9624 | 358.9511 | 18.0254 | 54.1576 | 1.0433 | 185.9197 | 6.0769 | 80.0262 | 3.2514 | 100.6795 | 2.5007 | 14.1191 | 0.29535 | 12.9171 | 0.39433 | 0.48116 | 0.03725 | 0.58999 | 0.054175 | 1.7193 | 0.081046 | 0.3471 | 0.031275 | 2.2653 | 0.16978 | 1.0427 | 0.082085 | 0.5227 | 0.042727 | 1.679 | 0.16554 | 1.8119 | 0.10388 | 0.31564 | 0.029029 | 2.5098 | 0.21758 | 0.53976 | 0.032398 | 1.557 | 0.14733 | 0.23772 | 0.01552 | 1.5798 | 0.13626 | 0.22825 | 0.020493 | 0.48628 | 0.045778 | 0.033727 | 0.0065167 | 3.8022 | 0.2082 | 3.8279 | 0.19865 | 0.030344 | 0.0083537 | 0.01819 | 0.0035857
157 | 1.349 | 0.030611 | 5.539 | 0.16211 | 7.6195 | 0.26234 | 9.0146 | 0.29472 | 8.9141 | 0.232 | 0.13669 | 0.003584 | 8.2247 | 0.26682 | 5293.2091 | 175.8018 | 324.3018 | 9.1228 | 368.5004 | 18.3707 | 54.4439 | 1.2486 | 181.0765 | 6.7189 | 81.1584 | 3.3564 | 102.4568 | 2.9666 | 13.8994 | 0.43819 | 13.1591 | 0.42763 | 0.46222 | 0.035863 | 0.56872 | 0.032917 | 1.7837 | 0.070983 | 0.34591 | 0.018384 | 2.3371 | 0.158 | 0.97398 | 0.078822 | 0.53271 | 0.037684 | 1.9387 | 0.13051 | 1.7951 | 0.11102 | 0.30588 | 0.020674 | 2.3032 | 0.14382 | 0.5096 | 0.031397 | 1.6666 | 0.10456 | 0.21898 | 0.016023 | 1.5174 | 0.098079 | 0.22682 | 0.023644 | 0.51787 | 0.04916 | 0.028005 | 0.0060223 | 3.6348 | 0.21467 | 3.8158 | 0.15071 | 0.025839 | 0.0059685 | 0.01322 | 0.0030342
158 | 1.3515 | 0.045279 | 5.5515 | 0.14608 | 7.6795 | 0.22161 | 9.0685 | 0.22645 | 8.785 | 0.2191 | 0.13672 | 0.0036928 | 8.1317 | 0.12905 | 5331.3494 | 113.2284 | 327.9133 | 15.3364 | 364.2178 | 15.3293 | 53.8085 | 1.0449 | 178.5482 | 5.8996 | 82.7001 | 4.8129 | 102.288 | 2.5248 | 13.966 | 0.41503 | 13.0865 | 0.4283 | 0.46967 | 0.036634 | 0.5437 | 0.03304 | 1.8024 | 0.090163 | 0.32142 | 0.023059 | 2.1131 | 0.15554 | 1.016 | 0.087022 | 0.46436 | 0.036624 | 1.9779 | 0.14776 | 1.7974 | 0.091206 | 0.33131 | 0.024637 | 2.2956 | 0.16451 | 0.50417 | 0.028066 | 1.5862 | 0.092577 | 0.22227 | 0.018445 | 1.6193 | 0.11282 | 0.23012 | 0.019544 | 0.53104 | 0.052867 | 0.037543 | 0.0069695 | 3.7658 | 0.17356 | 3.9843 | 0.15285 | 0.024001 | 0.005144 | 0.014096 | 0.0033937
207 | 1.3405 | 0.03479 | 5.5816 | 0.13246 | 7.8767 | 0.25284 | 9.2644 | 0.30902 | 8.9093 | 0.21672 | 0.13709 | 0.0025471 | 8.2428 | 0.13202 | 5482.4295 | 254.6051 | 330.1635 | 10.1055 | 346.1942 | 13.5422 | 55.196 | 1.2792 | 181.0727 | 5.6529 | 90.7068 | 15.9773 | 103.6188 | 2.7647 | 14.1798 | 0.34438 | 13.2653 | 0.54165 | 0.46432 | 0.040382 | 0.56444 | 0.033507 | 1.7495 | 0.072843 | 0.34725 | 0.022692 | 2.1538 | 0.12451 | 1.0466 | 0.11766 | 0.48702 | 0.034064 | 1.8362 | 0.1411 | 1.8429 | 0.10257 | 0.30052 | 0.021328 | 2.3556 | 0.12562 | 0.51025 | 0.024349 | 1.5705 | 0.1231 | 0.23585 | 0.01651 | 1.6558 | 0.10849 | 0.23473 | 0.021698 | 0.52326 | 0.055194 | 0.028582 | 0.0054033 | 3.7946 | 0.21322 | 3.9582 | 0.17877 | 0.026404 | 0.0064353 | 0.019136 | 0.0067449
208 | 1.3372 | 0.040609 | 5.5283 | 0.16633 | 7.6492 | 0.24907 | 8.9751 | 0.29316 | 8.7747 | 0.25921 | 0.13549 | 0.0042191 | 8.0867 | 0.18498 | 5308.0392 | 157.3884 | 320.3475 | 9.5549 | 340.2487 | 14.8381 | 54.1454 | 1.284 | 178.856 | 5.9635 | 80.574 | 3.3314 | 102.5202 | 3.0754 | 13.7679 | 0.43378 | 13.3086 | 0.40162 | 0.47354 | 0.039474 | 0.58061 | 0.044175 | 1.7702 | 0.073439 | 0.34475 | 0.026961 | 2.1746 | 0.15995 | 1.042 | 0.089408 | 0.49617 | 0.038034 | 1.7083 | 0.16341 | 1.7058 | 0.13839 | 0.31872 | 0.019136 | 2.5258 | 0.17827 | 0.50896 | 0.026773 | 1.6016 | 0.088813 | 0.22369 | 0.021346 | 1.5705 | 0.11397 | 0.22478 | 0.017654 | 0.51294 | 0.058155 | 0.029884 | 0.0070677 | 3.648 | 0.23797 | 3.822 | 0.3335 | 0.023705 | 0.0054652 | 0.016616 | 0.0045775
230 | 1.394 | 0.092798 | 5.4639 | 0.15779 | 7.8588 | 0.18046 | 9.0424 | 0.29518 | 9.1647 | 0.48587 | 0.13444 | 0.0034816 | 8.276 | 0.18665 | 5259.7066 | 135.8835 | 332.0618 | 9.5111 | 328.7284 | 13.8514 | 54.7628 | 1.0585 | 173.4408 | 5.4167 | 84.431 | 3.796 | 102.3114 | 2.9451 | 14.0581 | 0.38859 | 13.0022 | 0.411 | 0.4661 | 0.041896 | 0.57153 | 0.03664 | 1.8164 | 0.085546 | 0.35194 | 0.024484 | 2.1739 | 0.11729 | 1.1318 | 0.13578 | 0.49812 | 0.044428 | 1.7933 | 0.21596 | 1.8577 | 0.1465 | 0.32001 | 0.024343 | 2.4415 | 0.13609 | 0.51278 | 0.025364 | 1.5693 | 0.1035 | 0.22451 | 0.019861 | 1.5581 | 0.10669 | 0.24284 | 0.017243 | 0.52369 | 0.053353 | 0.026954 | 0.006692 | 4.1624 | 0.45048 | 3.9726 | 0.18377 | 0.026644 | 0.0068621 | 0.016526 | 0.0033582
231 | 1.3024 | 0.063329 | 5.3529 | 0.28792 | 7.5995 | 0.37469 | 8.7467 | 0.43851 | 8.6602 | 0.3919 | 0.12933 | 0.0064439 | 8.1573 | 0.35891 | 5060.8491 | 247.9092 | 319.8726 | 16.6966 | 319.8552 | 18.9632 | 52.5301 | 2.4978 | 166.0468 | 8.2503 | 78.2673 | 5.2631 | 108.2071 | 13.171 | 13.3198 | 0.67564 | 12.5336 | 0.69042 | 0.46548 | 0.03294 | 0.54307 | 0.036237 | 1.698 | 0.12502 | 0.33889 | 0.028522 | 2.1826 | 0.19037 | 0.9549 | 0.093323 | 0.46648 | 0.059657 | 1.7945 | 0.16763 | 1.7727 | 0.11699 | 0.31037 | 0.023295 | 2.2101 | 0.15931 | 0.48245 | 0.041738 | 1.5195 | 0.11104 | 0.21933 | 0.022235 | 1.507 | 0.12697 | 0.22267 | 0.019609 | 0.49341 | 0.047451 | 0.030594 | 0.0073054 | 4.0145 | 0.30819 | 5.9257 | 4.0275 | 0.029176 | 0.0079952 | 0.019414 | 0.003639
Average | 1.3452 |  | 5.5075 |  | 7.7057 |  | 9.0247 |  | 8.8533 |  | 0.13485 |  | 8.2326 |  | 5278.5194 |  | 325.1536 |  | 346.1064 |  | 54.1736 |  | 180.1978 |  | 82.7528 |  | 102.5246 |  | 13.9813 |  | 13.083 |  | 0.46594 |  | 0.55828 |  | 1.7762 |  | 0.34389 |  | 2.2049 |  | 1.024 |  | 0.49297 |  | 1.7975 |  | 1.8151 |  | 0.31861 |  | 2.3811 |  | 0.51476 |  | 1.5696 |  | 0.22462 |  | 1.5739 |  | 0.22781 |  | 0.51486 |  | 0.030957 |  | 3.8038 |  | 4.0036 |  | 0.026565 |  | 0.017581
2s | 0.044187 |  | 0.15046 |  | 0.19608 |  | 0.23972 |  | 0.27177 |  | 0.0047644 |  | 0.15729 |  | 195.7747 |  | 8.4818 |  | 28.0377 |  | 1.2819 |  | 12.4061 |  | 6.1942 |  | 4.0263 |  | 0.53282 |  | 0.42744 |  | 0.02038 |  | 0.037175 |  | 0.074131 |  | 0.022601 |  | 0.12602 |  | 0.11833 |  | 0.043262 |  | 0.18049 |  | 0.10671 |  | 0.021495 |  | 0.19243 |  | 0.031991 |  | 0.095089 |  | 0.014127 |  | 0.13407 |  | 0.016693 |  | 0.049476 |  | 0.0064034 |  | 0.29721 |  | 1.2302 |  | 0.0048763 |  | 0.0044679
GeoRem Preferred value | 1.37 |  | 5.67 |  | 8.81 |  | 9.51 |  | 9.51 |  | 0.15 |  | 8.08 |  | 5400 |  | 326 |  | 392 |  | 52 |  | 178 |  | 78 |  | 109 |  | 14.3 |  | 14 |  | 0.52 |  | 0.609 |  | 1.89 |  | 0.37 |  | 2.37 |  | 1.09 |  | 0.517 |  | 1.85 |  | 1.85 |  | 0.35 |  | 2.55 |  | 0.56 |  | 1.7 |  | 0.24 |  | 1.64 |  | 0.248 |  | 0.57 |  | 0.036 |  | 3.7 |  | 3.7 |  | 0.03 |  | 0.023
2s | 0.05 |  | 0.06 |  | 0.11 |  | 0.14 |  | 0.14 |  | 0.008 |  | 0.078 |  | 200 |  | 32 |  | 24 |  | 5 |  | 18 |  | 17 |  | 2 |  | 1.4 |  | 1.2 |  | 0.04 |  | 0.02 |  | 0.04 |  | 0.02 |  | 0.03 |  | 0.02 |  | 0.005 |  | 0.02 |  | 0.02 |  | 0.04 |  | 0.02 |  | 0.03 |  | 0.02 |  | 0.03 |  | 0.03 |  | 0.009 |  | 0.03 |  | 0.006 |  | 0.3 |  | 0.3 |  | 0.002 |  | 0.006

## Detection Limits
Table 4h. Average detection limits for LA-ICP-MS trace element analysis.
Element | Detection limit (μg/g)
Na | 96.268
Mg | 2.3744
Al | 4.349
Si | 2252.4788
Ca (43) | 239.4672
Ca (44) | 151.5885
Ti | 0.61627
V | 0.09631
Cr | 2.2482
Mn | 1.9046
Fe | 10.0703
Co | 0.067222
Ni | 2.5342
Zn | 0.62247
Sr | 0.044927
Y | 0.015153
Zr | 0.020755
Nb | 0.0088667
La | 0.0060849
Ce | 0.0045485
Pr | 0.003124
Nd | 0.017603
Sm | 0.016902
Eu | 0.0058266
Gd (157) | 0.038152
Gd (158) | 0.015828
Tb | 0.0035055
Dy | 0.011086
Ho | 0.0024913
Er | 0.0071759
Tm | 0.0023165
Yb | 0.013544
Lu | 0.0029755
Hf | 0.007381
Ta | 0.0026819
Pb (206) | 0.009003
Pb (208) | 0.0065708
Th | 0.00029798
U | 0.00014583

## Internal Standards
Table 4i. Internal standardization values for LA-ICP-MS trace element analysis.
Sample | Si (wt. %) | Type
OM20-13 | 17.56 | Measured by EPMA
OM20-17 | 17.62 | Measured by EPMA
OM20-18 | 17.8 | Measured by EPMA
GSD-1G | 24.8675 | GeoREM
NISTSRM612 | 33.7021 | GeoREM
BCR-1 | 25.4892 | GeoREM
BIR-1G | 22.2032 | GeoREM