## METADATA
Sheet: VSMOW-SLAP calibration
This sheet contains raw isotope data of the three working standards (ICE, TAP, NOC) that have been analysed weekly during the study period. ICE and TAP were used for two-point VSMOW-SLAP calibration. NOC standard was used as quality control.
Variable | Description
Date | Date of the measurement
H2O | water mixing ratio measured by the Picarro L2140-i CRDS
d17O_raw | raw d17O value measured by the Picarro L2140-i CRDS; suffix ICE, TAP and NOC indicate the measured working standard
d18O_raw | raw d18O value measured by the Picarro L2140-i CRDS; suffix ICE, TAP and NOC indicate the measured working standard
d2H_raw | raw d2H value measured by the Picarro L2140-i CRDS; suffix ICE, TAP and NOC indicate the measured working standard
d17O_cal_NOC | VSMOW-SLAP calibrated d17O value of NOC; calibrated based on ICE and TAP data
d17Oprime_cal_NOC | calculated VSMOW-SLAP calibrated d'17O value of NOC
d18O_cal_NOC | VSMOW-SLAP calibrated d18O value of NOC; calibrated based on ICE and TAP data
d18Oprime_cal_NOC | calculated VSMOW-SLAP calibrated d'18O value of NOC
d2H_cal_NOC | VSMOW-SLAP calibrated d'18O value of NOC; calibrated based on ICE and TAP data
Exc17_cal_NOC | calculated VSMOW-SLAP calibrated 17O-excess of NOC
Exc2_cal_NOC | calculated VSMOW-SLAP calibrated d-excess of NOC
AV | Average over all calibration measurements
SD | Standard deviation over all calibration measurements
REF | VSMOW-SLAP calibrated reference values
Sheet: ICE_MixingRatioDependency
This sheet contains normalized isotope data (average and standard deviation (SD)) of the working standard ICE used for mixing ratio dependency calibration. The isotope data are expressed as difference relative to the value measured at a water mixing ratio of 10000 ppmv.
Sheet: TAP_MixingRatioDependency
This sheet contains normalized isotope data (average and standard deviation (SD)) of the working standard TAP used for mixing ratio dependency calibration. The isotope data are expressed as difference relative to the value measured at a water mixing ratio of 10000 ppmv.
Sheet: NOC_MixingRatioDependency
This sheet contains normalized isotope data (average and standard deviation (SD)) of the working standard NOC used for mixing ratio dependency calibration. The isotope data are expressed as difference relative to the value measured at a water mixing ratio of 10000 ppmv.
Sheet: MixingRatioDependencyFunction
This sheet contains the coefficients (mean and standard deviation (SD)) of the mixing ratio dependency functions for d17O, d18O and d2H of ICE, TAP and NOC. The mixing ratio dependency function has the form: Y = a/x+bx+c, where x = H2O and Y is d17O, d18O, or d2H.

## VSMOW-SLAP calibration
Date | H2O_ICE | d17O_raw_ICE | d18O_raw_ICE | d2H_raw_ICE | H2O_TAP | d17O_raw_TAP | d18O_raw_TAP | d2H_raw_TAP | H2O_NOC | d17O_raw_NOC | d18O_raw_NOC | d2H_raw_NOC | d17O_cal_NOC | d17Oprime_cal_NOC | d18O_cal_NOC | d18Oprime_cal_NOC | d2H_cal_NOC | Exc17_cal_NOC | Exc2_cal_NOC
2020-12-15 | 11030.7 | -14.2084 | -27.8215 | -200.8729 | 10803.6 | -4.4026 | -9.3536 | -51.0716 | 11260.6 | -8.8596 | -17.697 | -119.5166 | -8.949 | -8.9892 | -16.8652 | -17.0091 | -125.2738 | -8.4501 | 9.6481
2021-01-05 | 11349.2 | -14.2102 | -27.7796 | -201.0342 | 11187.4 | -4.3776 | -9.3052 | -51.1498 | 11238.4 | -8.8714 | -17.729 | -120.5194 | -8.9732 | -9.0138 | -16.9416 | -17.0867 | -126.1287 | 8.0512 | 9.4041
2021-01-12 | 11214.8 | -14.1844 | -27.7178 | -200.9312 | 11224.4 | -4.3742 | -9.2782 | -50.9876 | 11227.6 | -8.858 | -17.6948 | -120.2412 | -8.9735 | -9.014 | -16.9502 | -17.0955 | -125.9905 | 12.4118 | 9.6108
2021-01-19 | 11277.4 | -14.1366 | -27.628 | -200.6912 | 11304.8 | -4.3312 | -9.2102 | -51.0354 | 11237 | -8.8454 | -17.6334 | -120.2582 | -9.0057 | -9.0465 | -16.9665 | -17.1121 | -126.0891 | -11.3373 | 9.6431
2021-01-26 | 10308.6667 | -14.2457 | -27.8433 | -200.1743 | 11153.6 | -4.399 | -9.3218 | -51.4488 | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN
2021-02-02 | 11485 | -14.2446 | -27.8182 | -200.6844 | 11284 | -4.4362 | -9.3538 | -51.2138 | 11301.4 | -8.8918 | -17.7286 | -120.206 | -8.9464 | -8.9867 | -16.8978 | -17.0422 | -125.9489 | 11.5958 | 9.2332
2021-02-16 | 11295.4 | -14.2258 | -27.7598 | -200.6542 | 11337 | -4.3758 | -9.2704 | -50.9028 | 11369.8 | -8.8534 | -17.6192 | -119.8928 | -8.9495 | -8.9898 | -16.861 | -17.0048 | -125.8217 | -11.2383 | 9.0662
2021-02-23 | 10765.6 | -14.2124 | -27.756 | -200.7086 | 11025.6 | -4.4376 | -9.339 | -51.2616 | 10933.4 | -8.8956 | -17.7332 | -120.082 | -8.9639 | -9.0044 | -16.9382 | -17.0833 | -125.7933 | 15.6279 | 9.7123
2021-03-02 | 11013 | -14.225 | -27.7548 | -200.5658 | 11059.6 | -4.3974 | -9.2856 | -50.8792 | 11054.8 | -8.94 | -17.7738 | -119.988 | -9.0237 | -9.0646 | -17.0074 | -17.1537 | -125.9653 | -7.471 | 10.0942
2021-03-10 | 11210.6 | -14.2126 | -27.7398 | -200.5008 | 11171.4 | -4.3868 | -9.2924 | -51.1122 | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN | NaN
2021-03-17 | 11075.2 | -14.2528 | -27.8196 | -201.2462 | 11051.4 | -4.423 | -9.3322 | -51.0522 | 10871.4 | -8.907 | -17.7424 | -120.3372 | -8.9648 | -9.0053 | -16.9224 | -17.0672 | -125.9094 | 6.2047 | 9.4695
2021-03-23 | 11020.4 | -14.2356 | -27.8004 | -201.0864 | 11014.6 | -4.414 | -9.3216 | -50.9546 | 10946.2 | -8.8928 | -17.7098 | -120.2496 | -8.9634 | -9.0038 | -16.9045 | -17.049 | -125.9466 | -1.9143 | 9.2897
2021-03-30 | 11071.8 | -14.205 | -27.7308 | -200.3074 | 10963.6 | -4.3888 | -9.308 | -50.8358 | 10989 | -8.905 | -17.7392 | -120.107 | -9.0028 | -9.0436 | -16.9722 | -17.1178 | -126.2183 | -5.3402 | 9.5591
2021-04-06 | 11209 | -14.2636 | -27.8104 | -200.7384 | 11140.8 | -4.4448 | -9.3304 | -51.0524 | 11207.6 | -8.8922 | -17.7152 | -120.1888 | -8.9336 | -8.9738 | -16.9006 | -17.0451 | -125.9922 | 26.0276 | 9.213
2021-04-15 | 11125 | -14.2214 | -27.78 | -200.7914 | 11144 | -4.4012 | -9.323 | -50.83 | 11138.8 | -8.9004 | -17.737 | -119.9536 | -8.9842 | -9.0248 | -16.9398 | -17.0849 | -125.8573 | -3.9693 | 9.6608
2021-04-20 | 11181 | -14.2154 | -27.7542 | -200.67 | 11116 | -4.3696 | -9.2872 | -51.0292 | 11140.8 | -8.8926 | -17.708 | -120.1932 | -8.9961 | -9.0368 | -16.942 | -17.0871 | -126.039 | -14.7834 | 9.4967
2021-04-27 | 11273.6 | -14.2908 | -27.882 | -200.484 | 11298 | -4.444 | -9.3624 | -51.0312 | 11218.8 | -8.9376 | -17.7814 | -120.1466 | -8.9667 | -9.0071 | -16.9166 | -17.0613 | -126.076 | 1.2818 | 9.2569
2021-05-05 | 11229.6 | -14.2602 | -27.8206 | -200.618 | 11129.8 | -4.4558 | -9.3892 | -50.7362 | 11148.4 | -8.9484 | -17.8174 | -120.2132 | -8.9848 | -9.0254 | -16.9653 | -17.1109 | -126.2334 | 9.1213 | 9.4891
2021-05-12 | 11325.4 | -14.2702 | -27.8528 | -200.6228 | 11268.8 | -4.4202 | -9.3396 | -50.8088 | 11346.8 | -8.9052 | -17.743 | -119.9364 | -8.9568 | -8.9971 | -16.9041 | -17.0486 | -125.9267 | 4.5749 | 9.3064
2021-05-18 | NaN | NaN | NaN | NaN | 11051.6 | -4.4368 | -9.3828 | -50.9738 | 11106.4 | -8.9122 | -17.771 | -120.2724 | NaN | NaN | NaN | NaN | NaN | NaN | NaN
2021-05-26 | NaN | NaN | NaN | NaN | 11108.5 | -4.4612 | -9.4708 | -51.1005 | 11156.8333 | -8.9373 | -17.8432 | -120.0465 | NaN | NaN | NaN | NaN | NaN | NaN | NaN
2021-05-29 | 11078.6667 | -14.2638 | -27.8475 | -200.954 | 11126.1667 | -4.4339 | -9.3766 | -50.1598 | 11105.1667 | -8.9083 | -17.7456 | -119.3347 | -8.9553 | -8.9957 | -16.8891 | -17.0334 | -125.5385 | -2.0445 | 9.5746
2021-06-02 | 11173.8333 | -14.2363 | -27.8043 | -199.7253 | 11070.1667 | -4.4393 | -9.409 | -50.8375 | 11113.8333 | -8.9293 | -17.768 | -119.5258 | -8.9856 | -9.0262 | -16.9131 | -17.0578 | -125.9151 | -19.6934 | 9.39
2021-06-08 | 11233 | -14.233 | -27.7683 | -199.3933 | 11174.5 | -4.4245 | -9.3618 | -50.8545 | 11178.8333 | -8.8858 | -17.7128 | -119.2322 | -8.952 | -8.9923 | -16.9002 | -17.0446 | -125.7695 | 7.2275 | 9.4321
2021-06-15 | 11141.8333 | -14.3067 | -27.8782 | -199.798 | 11044.3333 | -4.5178 | -9.5097 | -51.0987 | 11166.3333 | -8.9827 | -17.8568 | -119.778 | -8.9644 | -9.0048 | -16.9135 | -17.0582 | -125.9909 | 1.9215 | 9.3171
2021-06-23 | 11132.3333 | -14.3005 | -27.8838 | -199.8883 | 11144.5 | -4.5122 | -9.509 | -50.9527 | 11156.8333 | -8.9373 | -17.8247 | -119.6417 | -8.9253 | -8.9653 | -16.8794 | -17.0235 | -125.8943 | 23.0616 | 9.1411
2021-06-29 | 11180.1667 | -14.2813 | -27.8477 | -199.945 | 11083.8333 | -4.5078 | -9.5003 | -51.0375 | 11160.5 | -8.9515 | -17.8112 | -119.8535 | -8.9503 | -8.9906 | -16.887 | -17.0312 | -126.0302 | 1.8754 | 9.0656
2021-07-07 | 11058.5 | -14.3087 | -27.9158 | -200.0922 | 11075.1667 | -4.5368 | -9.5372 | -51.0995 | 11142.6667 | -8.9778 | -17.8703 | -119.868 | -8.9484 | -8.9887 | -16.895 | -17.0394 | -125.9459 | 8.1193 | 9.2144
2021-07-14 | 11107.9167 | -14.313 | -27.9197 | -200.2572 | 11095.3333 | -4.526 | -9.5317 | -51.2352 | 11178.4167 | -8.9913 | -17.8821 | -119.9737 | -8.9656 | -9.006 | -16.9079 | -17.0525 | -125.9038 | -2.3162 | 9.3597
2021-07-24 | 11177.3333 | -14.3092 | -27.8877 | -200.32 | 11169.6667 | -4.4802 | -9.4635 | -50.9553 | 11158.1667 | -8.9582 | -17.8273 | -119.8683 | -8.9593 | -8.9997 | -16.905 | -17.0495 | -125.9196 | 2.4731 | 9.3201
2021-08-03 | 11197.3333 | -14.3005 | -27.8703 | -199.9145 | 11192.6667 | -4.5765 | -9.5637 | -51.4172 | 11156.3333 | -9.0055 | -17.9113 | -120.1123 | -8.9581 | -8.9985 | -16.9419 | -17.0871 | -126.0971 | 23.5043 | 9.4385
2021-08-10 | 11224.6667 | -14.2868 | -27.8378 | -199.7335 | 11144.8333 | -4.5198 | -9.4908 | -50.9758 | 11092.3333 | -9.0142 | -17.9092 | -119.9777 | -9.0036 | -9.0443 | -16.9938 | -17.1399 | -126.278 | 5.5235 | 9.6726
2021-08-18 | 11316.3333 | -14.2717 | -27.8225 | -199.7175 | 11311.8333 | -4.5147 | -9.485 | -50.8117 | 11164.3333 | -8.9908 | -17.8415 | -119.8005 | -8.9901 | -9.0307 | -16.9368 | -17.0818 | -126.1987 | -11.5234 | 9.2953
2021-08-24 | 11347.6667 | -14.3035 | -27.8495 | -199.513 | 11315 | -4.526 | -9.5015 | -50.9302 | 11283.1667 | -9.0073 | -17.879 | -119.7198 | -8.9859 | -9.0265 | -16.9528 | -17.0982 | -126.1505 | 1.3633 | 9.4723
2021-09-01 | 11066.5 | -14.291 | -27.8705 | -200.1228 | 11199.5 | -4.5442 | -9.5088 | -51.0802 | 11136.3333 | -9.0032 | -17.8625 | -119.931 | -8.9776 | -9.0182 | -16.923 | -17.0679 | -126.0034 | -6.3426 | 9.3808
2021-09-08 | 11186.3333 | -14.3088 | -27.8415 | -199.6245 | 11202.1667 | -4.5462 | -9.4918 | -50.9313 | 11188.5 | -9.0032 | -17.8557 | -119.6553 | -8.9684 | -9.0089 | -16.9385 | -17.0836 | -126.037 | 11.2543 | 9.4712
2021-09-17 | 11248.1667 | -14.3103 | -27.8797 | -199.6708 | 11323.8333 | -4.5553 | -9.5503 | -50.7978 | 11311 | -9.039 | -17.9415 | -119.7713 | -8.9984 | -9.0392 | -16.9749 | -17.1206 | -126.1986 | 0.51473 | 9.6006
2021-09-21 | 11168.3333 | -14.335 | -27.9017 | -199.8122 | 11225 | -4.557 | -9.5373 | -51.0087 | 11166.1667 | -9.0638 | -17.9857 | -119.802 | -9.0109 | -9.0518 | -17.0157 | -17.1621 | -126.0549 | 9.8197 | 10.0707
2021-09-29 | 11232.8333 | -14.3362 | -27.866 | -199.2893 | 11147.5 | -4.5525 | -9.5097 | -50.499 | 11129.3333 | -9.0192 | -17.8992 | -119.655 | -8.9685 | -9.009 | -16.961 | -17.1065 | -126.413 | 23.2318 | 9.2748
2021-10-13 | 11017.6667 | -14.3347 | -27.899 | -200.1447 | 11034.8333 | -4.539 | -9.4735 | -50.4137 | 11025 | -9.0302 | -17.9008 | -119.8035 | -8.9874 | -9.028 | -16.9671 | -17.1127 | -126.2167 | 7.5267 | 9.5203
2021-10-27 | 11015.6667 | -14.3055 | -27.872 | -199.8275 | 10943.5 | -4.5475 | -9.4932 | -51.0353 | 10937.6667 | -9.0288 | -17.9015 | -119.8197 | -8.9948 | -9.0354 | -16.9694 | -17.1151 | -126.0512 | 1.3146 | 9.7044
2021-11-09 | 11071.1667 | -14.0743 | -27.7685 | -200.0008 | 11087.6 | -4.2798 | -9.3476 | -50.7584 | 11047.3333 | -8.738 | -17.702 | -119.7355 | -8.9552 | -8.9956 | -16.8971 | -17.0415 | -126.0363 | 2.3466 | 9.1405
2021-11-16 | 11064.1667 | -14.0755 | -27.7753 | -200.2545 | 10946.5 | -4.2947 | -9.3995 | -50.8585 | 10998.6667 | -8.8012 | -17.7982 | -119.8378 | -9.0093 | -9.0501 | -16.9612 | -17.1067 | -125.9698 | -17.7927 | 9.72
2021-11-26 | 11202.6667 | -14.119 | -27.8193 | -200.3375 | 11225.5 | -4.3052 | -9.402 | -50.7508 | 11369.6667 | -8.7973 | -17.7875 | -119.6993 | -8.9801 | -9.0207 | -16.9295 | -17.0744 | -125.855 | -5.3956 | 9.5807
2021-11-30 | 11185 | -14.0732 | -27.7548 | -200.1263 | 11138.6667 | -4.2985 | -9.3685 | -50.6047 | 11478.1667 | -8.7872 | -17.7902 | -119.691 | -8.9944 | -9.0351 | -16.9793 | -17.1251 | -126.0171 | 6.9035 | 9.8169
2021-12-09 | 11139.3333 | -14.0938 | -27.8043 | -200.566 | 11082.1667 | -4.2702 | -9.3447 | -50.923 | 11246 | -8.7725 | -17.7668 | -119.9085 | -8.9857 | -9.0263 | -16.9466 | -17.0918 | -125.8656 | -1.8259 | 9.7073
2021-12-15 | 11011.67 | -14.08 | -27.77 | -200.68 | 11112.33 | -4.26 | -9.34 | -50.84 | 11064.5 | -8.76 | -17.74 | -120.16 | -8.9851 | -9.0257 | -16.9381 | -17.0832 | -126.1007 | -5.7557 | 9.404
2021-12-21 | 10934.5 | -14.0857 | -27.8158 | -200.622 | 10851.6667 | -4.2857 | -9.3585 | -51.1233 | 10786.8333 | -8.7757 | -17.7647 | -120.19 | -8.9842 | -9.0248 | -16.9319 | -17.0769 | -126.0084 | -8.2555 | 9.4466
2022-01-05 | 11269 | -14.1412 | -27.876 | -200.1357 | 11417.5 | -4.3403 | -9.4357 | -50.7828 | 11363.5 | -8.8015 | -17.8278 | -119.7933 | -8.9553 | -8.9957 | -16.9257 | -17.0706 | -126.0193 | 17.604 | 9.3863
AV | 11141.3607 | -14.234 | -27.8195 | -200.2947 | 11137.2401 | -4.4341 | -9.4012 | -50.9437 | 11153.1833 | -8.915 | -17.7932 | -119.9253 | -8.9742 | -9.0147 | -16.9327 | -17.0777 | -125.9931 | 2.3341 | 9.4683
SD | 176.0431 | 0.078062 | 0.059997 | 0.48383 | 123.6473 | 0.088625 | 0.091129 | 0.22803 | 136.22 | 0.082778 | 0.081461 | 0.26621 | 0.021926 | 0.022124 | 0.035725 | 0.03634 | 0.1873 | 10.9503 | 0.22908
REF | – | -14.2404 | -26.8485 | -203.76 | – | -4.5397 | -8.6382 | -59.2434 | – |  |  |  | -8.9619 | -9.0023 | -16.9109 | -17.0555 | -125.6938 | 2.9634 | 9.5931

## ICE_MixingRatioDependency
sequence | STD | d17O | d17Oprime | d18O | d18Oprime | d2H | Exc17 | Exc2 | H2O | d17OSD | d18OSD | d2HSD | Exc17SD | Exc2SD | H2OSD
2021-05-26 | ICE | -0.0050726 | -0.0050726 | 0.12475 | 0.12474 | -0.53351 | -70.9371 | -1.5315 | 30073.6667 | 0.012964 | 0.018974 | 0.15297 | 2.8475 | 0.30476 | 134.8217
2021-05-26 | ICE | -0.0091305 | -0.0091305 | 0.088834 | 0.08883 | -0.46876 | -56.0329 | -1.1794 | 26212 | 0.0075425 | 0.0018856 | 0.052444 | 6.6275 | 0.067529 | 379.7163
2021-05-26 | ICE | -0.0042272 | -0.0042272 | 0.058403 | 0.058402 | -0.33926 | -35.0632 | -0.80649 | 21888.25 | 0.012492 | 0.0095459 | 0.03288 | 7.489 | 0.043487 | 127.1614
2021-05-26 | ICE | -0.031111 | -0.031111 | -0.0048587 | -0.0048587 | -0.44034 | -28.5457 | -0.40147 | 17040.0833 | 0.013435 | 0.014496 | 0.053033 | 5.7571 | 0.169 | 3.4177
2021-05-26 | ICE | -0.01192 | -0.01192 | 0.0037134 | 0.0037133 | -0.30959 | -13.881 | -0.3393 | 14114.4167 | 0.016853 | 0.01426 | 0.077192 | 9.3516 | 0.19127 | 44.9013
2021-05-26 | ICE | 0.00016884 | 0.00016884 | 0.0049992 | 0.0049992 | -0.13701 | -2.4707 | -0.177 | 11078.6667 | 0.011785 | 0.014142 | 0.014849 | 4.2751 | 0.12799 | 29.2271
2021-05-26 | ICE | -0.00025385 | -0.00025385 | -0.0075161 | -0.0075161 | 0.20599 | 3.7147 | 0.26612 | 8378.25 | 0.010253 | 0.035827 | 0.092395 | 9.0572 | 0.37901 | 4.1248
2021-05-26 | ICE | 0.044636 | 0.044635 | 0.0057707 | 0.0057706 | 0.46107 | 41.5886 | 0.41491 | 6848.4167 | 0.011314 | 0.0095459 | 0.07295 | 6.2922 | 0.14932 | 28.1664
2021-05-26 | ICE | 0.071351 | 0.071348 | 0.038602 | 0.038601 | 0.99082 | 50.967 | 0.68201 | 5308.0833 | 0.004714 | 0.002357 | 0.053033 | 6.0617 | 0.071889 | 20.8597
2021-05-26 | ICE | 0.1274 | 0.12739 | 0.062175 | 0.062173 | 1.7527 | 94.5649 | 1.2553 | 3815.0833 | 0.013081 | 0.019445 | 0.026163 | 2.7085 | 0.1294 | 16.1456
2021-05-26 | ICE | 0.39387 | 0.39379 | 0.3574 | 0.35733 | 3.671 | 205.1181 | 0.81181 | 2159.5 | 0.027224 | 0.03288 | 0.0042426 | 9.7554 | 0.2588 | 11.078
2021-10-20 | ICE | 0.0084314 | 0.0084313 | 0.12146 | 0.12145 | -0.35321 | -55.6954 | -1.3249 | 29597.8333 | 0.0071889 | 0.039244 | 0.046551 | 14.0181 | 0.36051 | 46.669
2021-10-20 | ICE | 0.010883 | 0.010883 | 0.0828 | 0.082797 | -0.35313 | -32.8338 | -1.0155 | 25329.25 | 0.0049497 | 0 | 0.091924 | 5.0213 | 0.091924 | 103.8268
2021-10-20 | ICE | 0.0090232 | 0.0090231 | 0.065828 | 0.065826 | -0.32054 | -25.7329 | -0.84717 | 21161.5 | 0.00094281 | 0.0056569 | 0.10312 | 4.029 | 0.057865 | 128.9291
2021-10-20 | ICE | -0.011689 | -0.01169 | 0.030855 | 0.030854 | -0.30171 | -27.9805 | -0.54855 | 16818.9167 | 0.01096 | 0.0080139 | 0.017796 | 6.7669 | 0.046315 | 124.0972
2021-10-20 | ICE | 0.0089386 | 0.0089386 | 0.033683 | 0.033683 | -0.33504 | -8.8459 | -0.60451 | 13905 | 0.00011785 | 0.0097816 | 0.21225 | 5.4317 | 0.2905 | 158.8633
2021-10-20 | ICE | -0.0078851 | -0.0078851 | -0.00077563 | -0.00077563 | -0.068792 | -7.4756 | -0.062587 | 10929.3333 | 0.0032998 | 0.021802 | 0.051383 | 8.4934 | 0.2258 | 95.9308
2021-10-20 | ICE | 0.014772 | 0.014772 | 0.0014531 | 0.0014531 | 0.12887 | 14.0047 | 0.11725 | 8259 | 0.004714 | 0.0095459 | 0.25503 | 0.40223 | 0.3314 | 78.4889
2021-10-20 | ICE | 0.043939 | 0.043938 | 0.04157 | 0.041569 | 0.41537 | 21.9895 | 0.082818 | 6752.3333 | 0.0062461 | 0.0095459 | 0.22934 | 1.152 | 0.30571 | 46.669
2021-10-20 | ICE | 0.05104 | 0.051039 | 0.050827 | 0.050826 | 0.89046 | 24.2028 | 0.48384 | 5163 | 0.005539 | 0.018267 | 0.041837 | 15.5393 | 0.1043 | 19.3276
2021-10-20 | ICE | 0.10582 | 0.10582 | 0.094544 | 0.094539 | 1.4632 | 55.9006 | 0.70686 | 3683.25 | 0.030523 | 0.040187 | 0.037948 | 9.1387 | 0.35945 | 13.3172
2021-10-20 | ICE | 0.3548 | 0.35473 | 0.4063 | 0.40622 | 3.5353 | 140.2494 | 0.28486 | 2015 | 0.0025927 | 0.0098995 | 0.089449 | 2.7449 | 0.010253 | 8.0139
2022-01-05 | ICE | 0.017961 | 0.017961 | 0.12669 | 0.12668 | -0.25202 | -48.9288 | -1.2656 | 29951.8333 | 0.010017 | 0.0028284 | 0.16428 | 8.6247 | 0.14166 | 6.1283
2022-01-05 | ICE | -0.0059601 | -0.0059601 | 0.083062 | 0.083059 | -0.31502 | -49.8153 | -0.97952 | 25774.1667 | 0.002357 | 0.013081 | 0.061283 | 4.7129 | 0.16593 | 24.0416
2022-01-05 | ICE | -0.0043541 | -0.0043541 | 0.051861 | 0.05186 | -0.28727 | -31.7363 | -0.70216 | 21405.25 | 0.011196 | 0.0010607 | 0.088742 | 10.7801 | 0.080257 | 75.3069
2022-01-05 | ICE | -0.015681 | -0.015681 | 0.011574 | 0.011574 | -0.24935 | -21.792 | -0.34195 | 16966.3333 | 0.014024 | 0.035473 | 0.11337 | 5.0402 | 0.39716 | 122.0938
2022-01-05 | ICE | -0.01213 | -0.012131 | 0.0082315 | 0.0082315 | -0.14902 | -16.4768 | -0.21487 | 13912.25 | 0.0050676 | 0.013906 | 0.071418 | 2.4126 | 0.18267 | 125.04
2022-01-05 | ICE | -0.0027481 | -0.0027481 | -0.0037689 | -0.0037689 | -0.01752 | -0.75815 | 0.01263 | 10811.6667 | 0.014378 | 0.029227 | 0.044783 | 1.2898 | 0.2786 | 120.2082
2022-01-05 | ICE | 0.0061271 | 0.0061271 | 0.008403 | 0.0084029 | 0.039063 | 1.6903 | -0.028161 | 8190.3333 | 0.016146 | 0.02357 | 0.065643 | 3.5756 | 0.2542 | 124.9222
2022-01-05 | ICE | 0.051179 | 0.051178 | 0.028204 | 0.028203 | 0.26948 | 36.2868 | 0.043851 | 6549.1667 | 0.0016499 | 0.0076603 | 0.031584 | 5.8331 | 0.092867 | 61.9897
2022-01-05 | ICE | 0.10832 | 0.10831 | 0.089234 | 0.08923 | 0.56481 | 61.1995 | -0.14906 | 4880.75 | 0.022156 | 0.038773 | 0.037477 | 1.4144 | 0.27271 | 6.2461
2022-01-05 | ICE | 0.19428 | 0.19426 | 0.22518 | 0.22516 | 1.3226 | 75.3805 | -0.4788 | 3356.25 | 0.013081 | 0.01591 | 0.063404 | 21.9049 | 0.063875 | 10.4888
2022-01-05 | ICE | 0.54726 | 0.54711 | 0.62171 | 0.62152 | 2.7696 | 218.9521 | -2.204 | 1961.3333 | 0.065172 | 0.055979 | 0.103 | 35.6834 | 0.34483 | 27.5772

## TAP_MixingRatioDependency
sequence | STD | d17O | d17Oprime | d18O | d18Oprime | d2H | Exc17 | Exc2 | H2O | d17OSD | d18OSD | d2HSD | Exc17SD | Exc2SD | H2OSD
2021-05-26 | TAP | 0.069919 | 0.069916 | 0.24625 | 0.24622 | 0.33334 | -60.0878 | -1.6367 | 30046.75 | 0.0084853 | 0.024984 | 0.0002357 | 4.7906 | 0.19964 | 48.6725
2021-05-26 | TAP | 0.053094 | 0.053093 | 0.18526 | 0.18524 | 0.21551 | -44.716 | -1.2666 | 26172.6667 | 0.0053033 | 0.0095459 | 0.0032998 | 10.4136 | 0.073068 | 296.7491
2021-05-26 | TAP | 0.024718 | 0.024718 | 0.12604 | 0.12603 | 0.12109 | -41.8262 | -0.88722 | 22042.3333 | 0.012728 | 0.04938 | 0.049851 | 13.532 | 0.44489 | 215.9033
2021-05-26 | TAP | 0.039367 | 0.039366 | 0.097689 | 0.097684 | 0.084338 | -12.2115 | -0.69717 | 17059.5 | 0.017796 | 0.03677 | 0.022627 | 1.7224 | 0.31678 | 68.5894
2021-05-26 | TAP | 0.024886 | 0.024885 | 0.049907 | 0.049906 | 0.11926 | -1.4649 | -0.28 | 14150 | 0.014849 | 0.036062 | 0.059043 | 4.3055 | 0.34754 | 6.8354
2021-05-26 | TAP | 0.0030387 | 0.0030387 | 0.010958 | 0.010958 | -0.011745 | -2.7471 | -0.09941 | 11126.1667 | 0.019681 | 0.038302 | 0.15662 | 0.64628 | 0.14979 | 48.0833
2021-05-26 | TAP | -0.0040762 | -0.0040762 | -0.014699 | -0.0147 | 0.015755 | 3.6852 | 0.13335 | 8489.3333 | 0.0056569 | 0.016735 | 0.026281 | 3.2378 | 0.16016 | 25.6915
2021-05-26 | TAP | 0.02882 | 0.028819 | -0.01941 | -0.019411 | 0.086422 | 39.0681 | 0.2417 | 6934.5833 | 0.024867 | 0.032527 | 0.026988 | 7.6403 | 0.2872 | 94.1631
2021-05-26 | TAP | 0.03267 | 0.03267 | -0.024205 | -0.024206 | 0.29101 | 45.4502 | 0.48465 | 5303.6667 | 0.0057747 | 0.019445 | 0.13789 | 16.1647 | 0.29345 | 28.7557
2021-05-26 | TAP | 0.045142 | 0.045141 | -0.083512 | -0.083516 | 0.15526 | 89.2374 | 0.82335 | 3774.1667 | 0.008721 | 0.022863 | 0.053858 | 20.947 | 0.23676 | 31.5841
2021-05-26 | TAP | 0.16451 | 0.16449 | -0.0035111 | -0.0035111 | 0.74484 | 166.3454 | 0.77293 | 2180.4167 | 0.074246 | 0.049615 | 0.2258 | 48.1197 | 0.62273 | 4.8319
2021-10-20 | TAP | 0.062033 | 0.062031 | 0.21358 | 0.21355 | 0.3597 | -50.7258 | -1.3489 | 29849.6667 | 0.004714 | 0.0073068 | 0.34106 | 0.84181 | 0.28261 | 28.9914
2021-10-20 | TAP | 0.026708 | 0.026708 | 0.17034 | 0.17032 | 0.30487 | -63.2228 | -1.0578 | 25577.1667 | 0.0061283 | 0.0091924 | 0.38938 | 1.2568 | 0.31584 | 23.8059
2021-10-20 | TAP | 0.039767 | 0.039766 | 0.13846 | 0.13845 | 0.1977 | -33.3333 | -0.90994 | 21349.6667 | 0.013435 | 0.0038891 | 0.34271 | 11.4224 | 0.37382 | 127.5149
2021-10-20 | TAP | 0.033823 | 0.033823 | 0.10312 | 0.10312 | 0.18012 | -20.6233 | -0.64486 | 17060.1667 | 0.0017678 | 0.0050676 | 0.22333 | 0.92527 | 0.18279 | 64.111
2021-10-20 | TAP | 0.0096319 | 0.0096318 | 0.056771 | 0.056769 | 0.1257 | -20.3423 | -0.32846 | 14116.4167 | 0.019328 | 0.019799 | 0.27389 | 8.8622 | 0.11549 | 67.293
2021-10-20 | TAP | 0.0074555 | 0.0074555 | 0.021018 | 0.021018 | 0.067538 | -3.642 | -0.10061 | 11051.75 | 0.024984 | 0.029345 | 0.18173 | 9.4554 | 0.053033 | 16.3813
2021-10-20 | TAP | -0.011881 | -0.011881 | -0.033494 | -0.033495 | -0.10763 | 5.804 | 0.16032 | 8323.9167 | 0.017324 | 0.037595 | 0.18008 | 2.637 | 0.12068 | 54.3294
2021-10-20 | TAP | -0.039421 | -0.039422 | -0.051244 | -0.051245 | -0.26705 | -12.3641 | 0.14291 | 6848.5 | 0.0049497 | 0.039834 | 0.1578 | 16.2599 | 0.16087 | 0.70711
2021-10-20 | TAP | -0.030632 | -0.030632 | -0.085398 | -0.085402 | -0.22338 | 14.4602 | 0.45981 | 5263.5833 | 0.030052 | 0.038891 | 0.066586 | 9.4577 | 0.24454 | 47.494
2021-10-20 | TAP | -0.026446 | -0.026447 | -0.13587 | -0.13588 | -0.22413 | 45.2991 | 0.86285 | 3734.5833 | 0.016853 | 0.034413 | 0.03182 | 1.416 | 0.30712 | 2.7106
2021-10-20 | TAP | 0.021351 | 0.021351 | -0.16624 | -0.16626 | -0.1463 | 109.1335 | 1.1836 | 2063.75 | 0.0098995 | 0.0048319 | 0.11809 | 12.5187 | 0.15674 | 2.9463
2022-01-05 | TAP | 0.035102 | 0.035101 | 0.17147 | 0.17146 | 0.37669 | -55.4279 | -0.99508 | 29597.8333 | 0.0071889 | 0.039244 | 0.046551 | 14.0181 | 0.36051 | 46.669
2022-01-05 | TAP | 0.030136 | 0.030136 | 0.13508 | 0.13507 | 0.34891 | -41.1805 | -0.73171 | 25329.25 | 0.0049497 | 0 | 0.091924 | 5.0213 | 0.091924 | 103.8268
2022-01-05 | TAP | 0.0091026 | 0.0091026 | 0.091563 | 0.091559 | 0.2693 | -39.2405 | -0.4632 | 21161.5 | 0.00094281 | 0.0056569 | 0.10312 | 4.029 | 0.057865 | 128.9291
2022-01-05 | TAP | 0.014905 | 0.014905 | 0.067114 | 0.067112 | 0.11224 | -20.53 | -0.42467 | 16818.9167 | 0.01096 | 0.0080139 | 0.017796 | 6.7669 | 0.046315 | 124.0972
2022-01-05 | TAP | 0.0074846 | 0.0074846 | 0.038739 | 0.038739 | 0.072856 | -12.9694 | -0.23706 | 13905 | 0.00011785 | 0.0097816 | 0.21225 | 5.4317 | 0.2905 | 158.8633
2022-01-05 | TAP | 0.00045472 | 0.00045472 | 0.011654 | 0.011654 | 0.048634 | -5.6988 | -0.044602 | 10929.3333 | 0.0032998 | 0.021802 | 0.051383 | 8.4934 | 0.2258 | 95.9308
2022-01-05 | TAP | -0.00077272 | -0.00077272 | -0.019804 | -0.019805 | -0.082644 | 9.6841 | 0.07579 | 8259 | 0.004714 | 0.0095459 | 0.25503 | 0.40223 | 0.3314 | 78.4889
2022-01-05 | TAP | 0.014236 | 0.014235 | -0.029113 | -0.029113 | -0.13614 | 29.6073 | 0.09676 | 6752.3333 | 0.0062461 | 0.0095459 | 0.22934 | 1.152 | 0.30571 | 46.669
2022-01-05 | TAP | 0.018476 | 0.018476 | -0.03444 | -0.034441 | -0.18303 | 36.6604 | 0.092489 | 5163 | 0.005539 | 0.018267 | 0.041837 | 15.5393 | 0.1043 | 19.3276
2022-01-05 | TAP | 0.07611 | 0.076107 | -0.0013552 | -0.0013552 | -0.19259 | 76.8224 | -0.18175 | 3683.25 | 0.030523 | 0.040187 | 0.037948 | 9.1387 | 0.35945 | 13.3172
2022-01-05 | TAP | 0.2195 | 0.21947 | 0.13732 | 0.13731 | 0.18697 | 146.9723 | -0.9116 | 2015 | 0.0025927 | 0.0098995 | 0.089449 | 2.7449 | 0.010253 | 8.0139

## NOC_MixingRatioDependency
sequence | STD | d17O | d17Oprime | d18O | d18Oprime | d2H | Exc17 | Exc2 | H2O | d17OSD | d18OSD | d2HSD | Exc17SD | Exc2SD | H2OSD
2021-05-26 | NOC | 0.010475 | 0.010475 | 0.15423 | 0.15421 | -0.21482 | -70.9501 | -1.4486 | 30048.9167 | 0.023217 | 0.03948 | 0.080728 | 2.2061 | 0.39657 | 0.35355
2021-05-26 | NOC | -3.5492e-05 | -3.5492e-05 | 0.10731 | 0.1073 | -0.1274 | -56.6916 | -0.98587 | 26362.5833 | 0.012021 | 0.014849 | 0.0068354 | 4.1475 | 0.12563 | 6.4818
2021-05-26 | NOC | -0.0011286 | -0.0011286 | 0.091274 | 0.09127 | -0.2124 | -49.3192 | -0.9426 | 22114.25 | 0.01591 | 0.015203 | 0.060104 | 7.8819 | 0.18173 | 21.8025
2021-05-26 | NOC | -0.040563 | -0.040564 | 0.018312 | 0.018312 | -0.18723 | -50.2325 | -0.33373 | 17097.0833 | 0.002357 | 0.0010607 | 0.071889 | 1.8084 | 0.063404 | 2.2392
2021-05-26 | NOC | -0.019627 | -0.019627 | 0.01178 | 0.011779 | -0.20115 | -25.8463 | -0.29539 | 14154.1667 | 0.010489 | 0.0035355 | 0.095577 | 8.6825 | 0.067293 | 6.5997
2021-05-26 | NOC | -0.0055849 | -0.0055849 | 0.010337 | 0.010337 | -0.2109 | -11.0429 | -0.2936 | 11105.1667 | 0.015321 | 0.015203 | 0.017206 | 7.2863 | 0.13883 | 16.2635
2021-05-26 | NOC | 0.007616 | 0.0076159 | -0.014097 | -0.014097 | 0.2876 | 15.059 | 0.40037 | 8492.9167 | 0.0071889 | 0.045844 | 0.63168 | 17.3908 | 0.99843 | 113.7263
2021-05-26 | NOC | 0.038474 | 0.038473 | 0.0195 | 0.0195 | 0.26302 | 28.1774 | 0.10702 | 6899.75 | 0.021213 | 0.011432 | 0.26151 | 15.258 | 0.17006 | 76.9568
2021-05-26 | NOC | 0.061008 | 0.061006 | 0.025439 | 0.025438 | 0.5236 | 47.5746 | 0.32009 | 5362.6667 | 0.0042426 | 0.0095459 | 0.067647 | 0.85068 | 0.008721 | 26.3987
2021-05-26 | NOC | 0.097163 | 0.097158 | 0.04334 | 0.043339 | 0.94427 | 74.2755 | 0.59755 | 3822.6667 | 0.010842 | 0.011785 | 0.089095 | 4.6036 | 0.18338 | 12.7279
2021-05-26 | NOC | 0.27088 | 0.27084 | 0.18494 | 0.18492 | 2.3564 | 173.2018 | 0.87693 | 2160.8333 | 0.0044783 | 0.028874 | 0.20318 | 11.0005 | 0.027813 | 5.1854
2021-10-20 | NOC | 0.019614 | 0.019614 | 0.13714 | 0.13713 | 0.0051345 | -52.7906 | -1.092 | 29592.0833 | 0.028638 | 0.045019 | 0.076957 | 4.6991 | 0.2832 | 61.8718
2021-10-20 | NOC | 0.030966 | 0.030965 | 0.14545 | 0.14544 | -0.024449 | -45.8286 | -1.1881 | 25375.3333 | 0.031113 | 0.05374 | 0.22934 | 2.5106 | 0.20058 | 74.7176
2021-10-20 | NOC | 0.0066643 | 0.0066643 | 0.074355 | 0.074352 | -0.0061155 | -32.5934 | -0.60095 | 21214.1667 | 0.0045962 | 0.0056569 | 0.043605 | 7.6786 | 0.0016499 | 58.9256
2021-10-20 | NOC | 0.023818 | 0.023818 | 0.051277 | 0.051276 | -0.10278 | -3.2556 | -0.513 | 16838.6667 | 0.0012964 | 0.00070711 | 0.069768 | 1.6882 | 0.075425 | 53.9758
2021-10-20 | NOC | 0.0059916 | 0.0059916 | 0.019376 | 0.019376 | -0.026532 | -4.2388 | -0.18154 | 13990.6667 | 0.019445 | 0.029463 | 0.070357 | 3.783 | 0.16535 | 16.4992
2021-10-20 | NOC | -0.007967 | -0.007967 | -0.0014956 | -0.0014956 | -0.028032 | -7.1773 | -0.016068 | 10978.25 | 0.0034177 | 0.0077782 | 0.0020035 | 0.73272 | 0.060222 | 8.3674
2021-10-20 | NOC | 0.013728 | 0.013728 | 0.0025769 | 0.0025769 | 0.048301 | 12.367 | 0.027686 | 8314.4167 | 0.023217 | 0.039834 | 0.025809 | 2.0133 | 0.29286 | 28.1664
2021-10-20 | NOC | 0.015662 | 0.015662 | 0.015219 | 0.015219 | 0.24497 | 7.6262 | 0.12322 | 6817 | 0.0061283 | 0.034295 | 0.08827 | 12.252 | 0.18609 | 64.111
2021-10-20 | NOC | 0.031807 | 0.031806 | -0.013883 | -0.013883 | 0.73938 | 39.1362 | 0.85045 | 5387.5 | 0.022627 | 0.0042426 | 0.38844 | 25.1124 | 0.42238 | 84.6171
2021-10-20 | NOC | 0.052576 | 0.052575 | 0.032951 | 0.03295 | 0.7408 | 35.1771 | 0.47719 | 3737.5 | 0.0041248 | 0.011314 | 0.0090745 | 1.92 | 0.099584 | 12.2565
2021-10-20 | NOC | 0.24741 | 0.24738 | 0.23055 | 0.23052 | 2.0393 | 125.6602 | 0.19489 | 2054.5 | 0.019563 | 0.02251 | 0.018738 | 7.6379 | 0.16134 | 8.0139
2022-01-05 | NOC | 0.013341 | 0.013341 | 0.12671 | 0.1267 | 0.098287 | -53.5557 | -0.91536 | 30030.4167 | 0.015674 | 0.015792 | 0.082496 | 7.3248 | 0.20883 | 207.0644
2022-01-05 | NOC | -0.014149 | -0.014149 | 0.076825 | 0.076822 | 0.064704 | -54.711 | -0.54989 | 25878.5833 | 0.015792 | 0.0018856 | 0.082142 | 16.9442 | 0.097227 | 148.6103
2022-01-05 | NOC | -0.016167 | -0.016167 | 0.047643 | 0.047641 | 0.026621 | -41.3214 | -0.35452 | 21574.25 | 0.00070711 | 0.021685 | 0.12044 | 10.9416 | 0.29392 | 21.8025
2022-01-05 | NOC | -0.013308 | -0.013308 | 0.027368 | 0.027367 | 0.0049538 | -27.7584 | -0.21399 | 17098.1667 | 0.0044783 | 0.029816 | 0.062225 | 11.5079 | 0.30076 | 65.9966
2022-01-05 | NOC | 0.0032528 | 0.0032528 | 0.019903 | 0.019902 | 0.014704 | -7.2557 | -0.14452 | 13973.0833 | 0.017088 | 0.0057747 | 0.043958 | 14.1348 | 0.090156 | 56.215
2022-01-05 | NOC | 0.00081484 | 0.00081484 | 0.0087897 | 0.0087897 | -0.037713 | -3.8261 | -0.10803 | 10875.5 | 0.0004714 | 0.016263 | 0.094281 | 8.2657 | 0.035827 | 32.7626
2022-01-05 | NOC | -0.0016231 | -0.0016231 | -0.017508 | -0.017508 | 0.075121 | 7.6213 | 0.21519 | 8256.0833 | 0.0062461 | 0.0091924 | 0.020506 | 11.242 | 0.053033 | 8.8388
2022-01-05 | NOC | 0.0094737 | 0.0094737 | -0.014539 | -0.014539 | 0.22887 | 17.1503 | 0.34518 | 6696.0833 | 0.02086 | 0.024631 | 0.063993 | 7.8039 | 0.13305 | 118.9118
2022-01-05 | NOC | 0.04495 | 0.044949 | -0.00088109 | -0.00088109 | 0.31712 | 45.414 | 0.32417 | 5065.3333 | 0.0090745 | 0.0056569 | 0.14661 | 6.1135 | 0.19186 | 33.234
2022-01-05 | NOC | 0.13658 | 0.13657 | 0.087174 | 0.08717 | 0.7342 | 90.5471 | 0.036811 | 3590.75 | 0.0083674 | 0.017206 | 0.043723 | 17.6879 | 0.18137 | 27.2236
2022-01-05 | NOC | 0.33111 | 0.33106 | 0.34006 | 0.34 | 1.5703 | 151.538 | -1.1502 | 2209.25 | 0.010017 | 0.0017678 | 0.017678 | 9.153 | 0.03182 | 16.3813

## MixingRatioDependencyFunction
coefficient | d17O mean | d17O SD | d18O mean | d18O SD | d2H mean | d2H SD
a_ICE | 937.0811 | 74.1099 | 1123.3932 | 101.2005 | 8886.9366 | 423.0333
b_ICE | 3.4649e-06 | 4.7183e-07 | 1.0094e-05 | 8.3247e-07 | 1.1235e-05 | 4.66e-06
c_ICE | -0.12595 | 0.013734 | -0.21526 | 0.021628 | -1.0144 | 0.11534
a_NOC | 676.6359 | 55.9864 | 666.6617 | 90.7721 | 5272.1562 | 672.2094
b_NOC | 2.9961e-06 | 7.0723e-07 | 9.4448e-06 | 1.0994e-06 | 1.4709e-05 | 5.9372e-06
c_NOC | -0.098365 | 0.015266 | -0.16369 | 0.024601 | -0.64159 | 0.1646
a_TAP | 294.7844 | 91.4122 | -26.8218 | 121.8352 | 161.7027 | 579.3464
b_TAP | 3.7465e-06 | 6.2237e-07 | 1.0463e-05 | 9.7896e-07 | 1.8785e-05 | 6.9819e-06
c_TAP | -0.065385 | 0.018545 | -0.099989 | 0.026029 | -0.20451 | 0.16064